PROGRAMMING FOR S2000-30KG-CA

1. Make sure scale is OFF. Remove the platter and find the CALIBRATION PORT located on the

front left side of the top cover.

2. The CALIBRATION PORT is covered by two brass SEALING BOLTS and a silver plate.

Remove the two SEALING BOLTS and the silver plate. You will see a red MICRO SWITCH.

3. Slide the MICRO SWITCH to the right. Put the platter back on the scale.

4. Turn scale ON and the WEIGHT display should flash “CAL” three times and then go blank.

5. Press the 4 key.

6. Press the following keys: 8 key, the MR key, and the C key.

7. Press the following keys: 3 key, the 9 key, and the C key. (60 lbs.) 2 + 0 For 30lbs.

8. Press the following keys: 2 key, the 0 key, and the C key.

9. Press the following keys: 9 key, the 2 key, and the C key.

10. Press the following keys: 0 key, the 7 key, and the C key.

11. Press the following keys: 0 key, the M3 key, and the C key.

12. Press the following keys: 0 key, the 8 key, and the C key.

IF last step doesn’t work 0 + M3, and the C key.

13. Once the scale displays go blank, press the 0 key. The scale will display Err 2

14. Turn scale OFF, remove the platter, slide the MICRO SWITCH to the left, replace SEALING

BOLT and silver plate, and replace the platter. You are done.