| S-12 ELECTRONIC WEIGHING SCALE |  | TROUBLE SHOOTING PROCEDURE |
| --- | --- | --- |
|  |  | P.1 |
| ouble Causes | Inspection Steps | Trouble-shooting Procedures |
| is no display after | 1. Check power supply | .The power for AC current is on. |
| ing switch on. | 2. Check poor contact. | Connect tightly from the plus to the wall receptacle. |
|  | 3. Check ON/OFF switch for poor | Press switch serveral times to check smooth connection |
|  | contact. |  |
|  | 4. Check fuse-holder | .Tight fuse-holder by turning in clockwise direction. |
|  | 5. Check internal circuit. | Inspect and repair. |
| •lay unstable. | 1.There is a strong wind. | .Replace to a vindless.vibration free place. |
|  | 2.There is a high-frequency | .Replace to a place without interference. |
|  | interference. |  |
|  | 3.Scale platform is not level. | Replace to a level place: may use adjustment legs. |
| key can not operate | 1.Check key function. | .Refers to damaged keys if it does not operate. |
| specified function | 2.Trouble in the interior wires | Check and repair |
| weight. | 1.There is an adjustment port | .The value displayed will increase vhile turning in |
|  | at the right side on the | clockwise direction. and reduce while turning in |
|  | body shell of the scale as | counterclockwise direction. Adjust by using standard |
|  | the figure follow= | weights. Make a microadjustment with a small flat head |
|  |  | screwdriver till the value of weight is accurate. |
|  | 2.Put on the same weight objects | The cause of this condition refers to a protective |
|  | every time to check the weight | device of Load Cell is contacted by Load Cell. |
|  | displayed is kept or reduced. |  |
| e is not zero value | 1.Check the O/P signal wires of | Proper weld the four wires at J3 on the I/P PC board: |
| lay after turning | LOAD CELL for poor contact. | the ,2th (yellow) ,3th (black) from |
| switch. |  | the left to right. |
|  | 2. Check LOAD CELL for proper | .Using resistance meter measure the resistance between |
|  | resistance. | the Ith and 2th wire. approximately 350 ohm :between |
|  |  | the 3 th and 4th wire, approximately 400 ohm. If the |
|  |  | resistance measured is not proper. refers to the |
|  |  | deviation of LOAD CELL. |
| is no display after | 1. Check U11 and U12 for | Measure with a voltmeter. The output voltage on the 3r |
| ing on the switch. | proper output voltage. | leg of U11 is 8V: the 3rd leg of U12 is +5V. |