1.5HARDY

HI2151/20WC

Code d’erreur.

| ERROR MESSAGE | DEFINITION |
| --- | --- |
| - - HI - - | Load cell/point signal represents a weight higher than scale capacity. |
| Error | Invalid character entry. |
| Err 1 | Invalid grad size. Grad size must be 1, 2, 5, 10, 20, 50, 100, 200, 500. |
| Err 2 | NTEP grad size error. Grad size must be greater than the full scale divided by the range |
| Err 3 | Average out of range. Average must be from 1 to 200. |
| Err 4 | Instrument in motion when value entered. |
| Err 5 | Not in Gross mode when trying to enter CAL menu. |
| Err 6 | Invalid zero. Zero must be less than zero tolerance. |
| Err 7 | # Invalid grad size. Grad size must be greater than the full scale divided by the range minus one valid grad size. |
| Err 8 | Unable to enter Calibration Menu without toggling ReCal segment on S3. |
| Err 9 | Invalid zero. When the instrument is in NTEP mode, the zero must be within the zero tolerance. |
| Err 10 | Invalid motion tolerance. Motion must be greater than 0 and greater then or equal to Grad Size. |
| Err 11 | Negative gross weight during acquire tare. (NTEP only) |
| Err 12 | All decades after decimal point must be active. (NTEP only) |
| Err13 | Full scale calibration of zero or negative value not allowed. |
| Err 14 | Scale capacity of zero or at negative value not allowed. |
| Err 15 | Zero tolerance of zero or a negative value not allowed. |
| Err 16 | Tare greater than span. (NTEP only) |
| Err 17 | Acquire tare cannot be used with zero gross weight. (NTEP only) |
| Err 18 | Difference between zero and span less then 100 counts. |
| Err 19 | Occurs when configuring Serial Port 1 and entering a negative address. |
| Err 20 | Option held when a print command was requested. Option hold active on remote functions TB when print command from serial port requested. (Serial Port only) |
| Err 24 | Clock malfunction |
| Err 25 | Invalid total decimal point. |
| Err 26 | Faulty time or date entry. |
| Err 40 | Software Cal sticker value too small. |
| Err 41 | Software Cal reference point weight out of range. |
| Err 42 | Software Cal load cell overloaded. |
| Err 43 | Software Cal range is zero error. |
| Err 44 | Sensitivity out of range. |
| Err 45 | Sensitivity out of range. |
| Err 50 | Too many BCD Options. |
| Err 52 | Too many Serial Options. |
| Err 54 | Invalid serial or BCD format entry. (____P or____) |
| Err 55 | Illegal analog weight span. Analog span weight minus analog zero weight must be larger than 2% of the calibrated span. |
| Err 60 | The number of load points found did not match the number entered. |
| Err 61 | No load points found. |
| Err 62 | Load point capacities or sensitivities do not match. |
| Err 63 | Check Sum Error on 1 or more load points. |
| Err 64 | Too many significant digits, after the decimal point, to be displayed |
| Err 97 | Checksum error. |
| Err 99 | Bad command sequence. (Serial Port only) |