Codes d'erreur

HB-8212

Error 01:	Function key has been inhibited via F10

Error 02:	The zero offset is out of the setting range

Error 03:	The system cannot tare because of a negative gross value is indicated.

Error 04:	The PRINT function cannot be proceeding because of a wrong action. That is, a weighing action is not coincident to the limitation, which has been set to the HB8212 via F12.

Error 05:	The PRINT function cannot be proceeded because of the weigher is

unstable

Error 06:	the system cannot tare or zero because of the system is in the data-

hold mode.

Error 07:	The sampling cannot proceed because of the system is in the data hold mode.

Error 11: 	The PRINT function cannot be proceeding because of the parallel printer error is detected.

Error 12:	The PRINT function cannot be proceeded because of the net weight value is negative.

Error 13:	The PRINT function cannot be proceeded because of a weighing action is not finish.

Error 14:	The index data has not been entered.

Error 15:	Double weighing data memory full.

Error 16:	Index data memory full.

Error 17:	Weighing data memory full.

Errfunc :	F-FUNCTION selected error.

ErrFCal:	The cal switch is not slid ON.

ErrFAna:	The analog O/P interface is not installed.

ErrF001:	The counting mode is not enabled.

ErrF002:	The counting mode is setting error.

ErrF003:	The sampling net weight is negative.

ErrF004:	The sampling net weight is under 5d.

ErrF005:	The sampling count cannot be entered as "0".

ErrF006:	The unit weight calculation is overflow.

ErrF091:	The data cannot be found to print list.

ErrF321:	Date setting error.

ErrF331:	Time setting error.

ErrF571:	The data cannot be found to print label.

ErrF621:	The minimum division setting error.

ErrF631:	The maximum capacity setting error. The maximum capacity should be a multiple of the division.

ErrF632:	The maximum capacity setting error. The resolution should be over or equal to 300.

ErrF633:	The maximum capacity setting error. The resolution should be under or equal to 10000.

ErrF641:	The HB8212 is not able to achieve zero calibration because of the input signal is unstable.

ErrF642:	The HB8212 is not able to achieve zero calibration because of the input signal is negative.

ErrF643:	The HB8212 is not able to achieve zero calibration because of the input signal is too large over (30mV).

ErrF651:	The HB8212 is not able to achieve span calibration because of the input signal is unstable.

ErrF652:	The HB8212 is not able to achieve span calibration because of the input signal is too small (under 0.6uV/d).

ErrF653:	The HB8212 is not able to achieve span calibration because of the input signal is too large (over 120uV/d).

ErrF654:	The balance weight is setting error.

1.	The balance weight must not be entered as a value equal to 0.

2. 	The balance weight must not be entered as a value greater than	 the 	maximum capacity.

3.	The balance weight must be entered as a value which is a multiple of the

minimum division.