Ishida

MT series

1.press the kg.lb/feed key to set the ratio of the span weight to the scale capacity.

For example, if the scale capacity is 300lb and the span weight is 60lb, then the ratio should be set to 60/300 which equal to 1/5. The display should therefore be set to LB5

2. Press TARE

3. Press ZERO

4. Place the span weight on the platter

5. Press TARE

6. Remove the span weigh from the platter

7. Repeat step 3 to 6 about two or three times.

END

The A/D initial value will be displayed