The procedure to enable "lb" unit is as follows:

1. Press M+ and switch on, display "unit" "out".

2. Press M+ and choose "out 1"(kg,lb,g) or choose "out 2"(kg, oz, lb, g).

3. Press ESC and off the scale.

Now switch on and you will have lb to choose from.

The question is lb or oz are not popular in other countries and we only design front panel with kg and g units. In that case, we will only make small lb and/or oz sticker to be adhered on panel.

JWG calibration procedure:

Press MR and switch on, display  "V 1.0" , "CAL" , "6K005".

Press Zero start zero calibration, when finish, the scale beep.

Put 2kg weigh mass and press zero for first calibration, when finish, the scale beep.

Put on 4kg weigh mass and press zero for second calibration, when finish, the scale beep

Put on 6kg weigh mass and then

Press zero for third calibration, when finish, the scale beep Press ESC, display PASS and of.

Regards,

4/1/2004.

Mr. Thomas Shen.