| CAPTER 1 GETTING STARTED |  |  |
| --- | --- | --- |
| 1-1 | Package contents --------------------------------------------------------------------- | 1 |
| 1-2 | Specifications -------------------------------------------------------------------------- | 1 |
| 1-3 | Preparation ----------------------------------------------------------------------------- | 2 |
| CHAPTER 2 SET UP |  |  |
| 2-1 | Calibration ------------------------------------------------------------------------------ | 2 |
| 2-2 | Power on setup ------------------------------------------------------------------------ | 4 |
| 2-3 | Offset value and keypad check ---------------------------------------------------- | 8 |
| 2-4 | Power supply voltage ---------------------------------------------------------------- | 10 |
| 2-5 | Recharging the battery --------------------------------------------------------------- | 10 |
| 2-6 | RS-232 serial port option ------------------------------------------------------------ | 11 |
| CHAPTER 3 MAINTENANCE |  |  |
| 3-1 | Replacing the main printed circuit board ---------------------------------------- | 12 |
| 3-2 | Replacing load cell -------------------------------------------------------------------- | 14 |
| CHAPTER 4 TROUBLESHOOTING |  |  |
| 4-1 | Preliminary checks -------------------------------------------------------------------- | 15 |
| 4-2 | Problems and solutions -------------------------------------------------------------- | 15 |
| CHAPTER 5 SPARE PARTS |  |  |
| 5-1 | LAC-3015/LAC-1260 ----------------------------------------------------------------- | 18 |
| 5-2 | LAW-3015/LAW-1260 ---------------------------------------------------------------- | 18 |

CHAPTER 1 GETTING STARTED

  - Package contents

- Scale body

- Weighing pan

- L-type hexagonal spanner (for disassembling delivery protector on the bottom

of scale)

- Sealing stamp (to seal scale after disassembling delivery protector)

- Power cord

- User manual

- Four (or, three) model labels

  - Specifications

LAC Series

| Model | LAC-3015 |  |  |  | LAC-1260 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Capacity | 30kg | 30kg | 15kg | 15kg | 12kg | 12kg | 6kg | 6kg |
| Display resolution | 2g | 5g | 1g | 2g | 1g | 2g | 0.5g | 1g |
| Counting resolution | 0.4g | 1g | 0.2g | 0.4g | 0.2g | 0.4g | 0.1g | 0.2g |
| Display | Liquid Crystal Display, with backlight, 14mm(H) X 7mm(W) WEIGHT --- 5 digits, UNIT WEIGHT --- 5 digits, TOTAL COUNT --- 5 digits |  |  |  |  |  |  |  |
| Pan size | 290 X 250 mm |  |  |  |  |  |  |  |
| Power | 110, 120, 220, 240 VAC +/- 10%, 50/60Hz, or by rechargeable battery |  |  |  |  |  |  |  |

LAW Series

| Model | LAW-3015 |  |  |  | LAW-1260 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Capacity | 30kg | 30kg | 15kg | 15kg | 12kg | 12kg | 6kg | 6kg |
| Display resolution | 2g | 5g | 1g | 2g | 1g | 2g | 0.5g | 1g |
| Counting resolution | 0.4g | 1g | 0.2g | 0.4g | 0.2g | 0.4g | 0.1g | 0.2g |
| Display | Liquid Crystal Display, with backlight, 14mm(H) X 7mm(W) WEIGHT --- 6 digits |  |  |  |  |  |  |  |
| Pan size | 290 X 250 mm |  |  |  |  |  |  |  |
| Power | 110, 120, 220, 240 VAC +/- 10%, 50/60Hz, or by rechargeable battery |  |  |  |  |  |  |  |

1-3 Preparation

Before using the scale, to release the delivey protector is must.  Follow the procedure below

to release.  After releasing, keep the protector plate and the screw, (13) & (12) respectively.

Remember to install them back whenever afterwards to deliver is needed.

- At the bottom of the scale body, loosen and remove the hexagonal screw. --- (12)

- Remove the protector plate, (13), followed by to seal the plate place, (8), by the sealing

stamp enclosed in the package.

CHAPTER 2 SET UP

2-1 Calibration

Note！: The regulations and/or metrological controls governing to conduct calibration vary from country to country. Please do follow the local metrological controls and/or regulations imposed by the government.

Calibration is done by software to reduce the need to repeatedly load and unload weight.

Calibration can be done by the following procedures:

LAC Series

- Turn off scale and leave the pan empty.

- Keep SMPL key pressed and turn on the scale, 0 will be displayed in UNIT WEIGHT and

TOTAL COUNT displays.

- Press number 11, then press SMPL, the words SET ZERO 0 will be shown in the display

and word ZERO will blink.          SET  ZERO  0

- When the display shows SPAN 0 -g- , zero calibration is completed and scale is now

waiting for span calibration.            SPAn  0  -g-

- Put on the weight desired for calibration and key in value of the weight (any weight within

the set capacity of the scale can be set as calibration weight.)

Press CE if you want to clear the set weight.            SPAn 5000g -g-

- Press SMPL, then span calibration will be executed.

(1). If span calibration is completed, the display will show SPAn PASS -g-.  Remove

calibration weight from the pan, then press SMPL twice, scale will start power-on

self-test and then will be ready for use.

(2). If the alarm sounds TWO beeps, it indicates that there is no value has been entered.

Enter the correct value and press SMPL again to complete calibration.

(3). If the alarm sounds THREE beeps, it indicates that the measured weight is incorrect (not within the acceptable range.)  Check the calibration weight and put on new weight(s) with the correct range.

LAW Series

(Press ZERO at any step in the procedure for leaving the current step and to enter normal

weighing mode, which will result in the uncompleted steps to be ignored.)

- Turn off the scale and leave the pan empty.

- Keep MC/SET key pressed and turn on the scale, CAL will be displayed.

- Press M+/MR, display will show CAL zro.  Zero calibration will then be executed.

CAL zro

- After zero calibration is completed, display will show the weight should be put on for span

calibration.  For each model of LAW, three different weights could be selected for

calibration.  Press MODE to select one the weights put for calibration.  Weights could be

used is as following:

| Model | LAW-1260 (6kg) | LAW-1260 (12kg) | LAW-3015 (15kg) | LAW-3015 (30kg) |
| --- | --- | --- | --- | --- |
| One of the weights could be used (kg) | 1 | 2 | 2 | 5 |
|  | 2 | 5 | 5 | 10 |
|  | 5 | 10 | 10 | 20 |

- Put on the weight as that shown in display, display will show CAL.  PASS will be displayed when calibration is completed.  PASS

- Remove calibration weight from the pan, then press ZERO to enter normal weighing mode.  Or, to enter the POWER ON SETUP mode stated in the succeeding Section.

2-2 Power on setup

Power On Setup function is provided to set up functions including capacity-resolution, level of

digital filtering, auto power-off, backlight, zero-band, operation mode (LAW only), and

power-on operation mode (LAW only).

2-2-1 LAC Series

Follow the steps below to set up each function:

- Turn off the scale.

- Keep ZERO pressed and turn on the scale.  Display will show  SEt  SCALE  Func..

Now, you can press numbers 1, 2, 3, 4, or 0 respectively to set up each function.

CAPACITY-RESOLUTION

- Press ” 1 ”, SEt 1. will be shown in WEIGHT display, current setup of

capacity and display resolution will be shown respectively in

#### SEt 1.  15Kg  1g

UNIT WEIGHT and TOTAL COUNT displays.

Press “ 1 ” repeatedly until the desired

SEt 1.  15Kg  1g

capacity-resolution is displayed.

(The available capacity-resolution setup is shown in Section 1-2.)

3.  Press ZERO to complete Power On Setup (scale will start power-on self-test and then will

be ready for use.)  Or, just press other numbers (2, 3, 4, or 0) to set up other functions.

DIGITAL FILTERING

This function is provided for weighing vibrating objects, or to prevent the scale from vibration

that may affect stability of the scale.

- Press “ 2 “,  SEt 2.  and  Filt  will be shown in WEIGHT and UNIT WEIGHT displays,

current setup of digital filtering level will be shown in TOTAL display.

Press “ 2 “ repeatedly to choose

### SEt 2.  Filt  -2-

digital filtering level as 1 or 2.

Level 1:	Reaction is faster; effect to filter vibration is less

Level 2:	Reaction is slower; effect to filter vibration is more

AUTO POWER-OFF

When auto power-off is enabled, scale will be automatically switched off after there has been

no load on the pan (and, scale is stable at zero) for the time duration set.

- Press “ 3 “,  SEt 3.  and  OFFt  will be shown in WEIGHT and UNIT WEIGHT

displays , current setup of auto power off time will be shown in TOTAL COUNT display.

Five options, 5, 10, 30, 60 minutes, and OFF (disable) are available.

Press “ 3 “ repeatedly to choose the

SEt 3.  OFFT  30

desired setup.  When OFF is chosen,

scale will always be ON after it has been powered on.

BACKLIGHT

Backlight is provided for having better visibility when the scale is used in dark environment.

- Press “ 0 “,  SEt 0.  and  Light  will be shown in WEIGHT and UNIT WEIGHT

displays, current setup of backlight function will be shown in TOTAL COUNT display.

- Press “ 0 “ repeatedly to                 Set 0. Light ONOFF

choose setup as ON, ONOFF Or OFF.

ON:        Backlight will be always ON.

ONOFF:    Backlight will be ON only when there is load on the pan, and it is greater

than 9 times of the set display resolution.

OFF:	     Backlight will be always OFF.

ZERO BAND

Zero-band function is used to define the weight range within which the scale will remain zero

indication whenever there is disturbance.  Weight range is expressed in terms of number of

display divisions, both for positive and negative directions.

- Press “ 4 “,  SEt 4.  and  ZERO  will be shown in WEIGHT and UNIT WEIGHT

displays, current setup of zero band will be shown in TOTAL COUNT display.

Press “ 4 “ repeatedly until the desired number

d 1

of display division is shown.  There are six choices;

0d (d0), 1d (d1), 2d (d2), 3d (d3), 4d (d4), and 5d (d5).

2-2-2 LAW Series

Follow the steps below to set up each function:

- Turn off the scale.

- Keep MC/SET pressed and turn on the scale.  Display will show  CAL

- Now, you can press MC/SET to enter the power-on setup mode.

- Press MC/SET again to go to the next setup when current setup is completed.

- Press ZERO at any stage to go to normal weighing mode.

CAPACITY – RESOLUTION

- Press MC/SET right after display shows CAL.  Display will show the current

capacity-resolution setup.

- Press M+/MR 1 repeatedly until the desired capacity-resolution is displayed. (The

available capacity-resolution setup is shown in Section 1-2.)

OPERATION MODE SELECTION

This function is provided for you to enable those weighing units you want to have in normal

operation.  Those disabled units will not be shown when you press MODE in normal

operation.

- Press MC/SET until display shows  on 	g.

- Press M+/MR to view the status of the six available units.  They are:

kg (or	g)		lb (or	oz)		tl.T (TAEL, Taiwan)

tl.H (TAEL, Hong Kong) 	pcs (piece counting) 	% (percentage)

- Press MODE to enable (ON) or to disable (OFF) each of it.

POWER-ON OPERATION MODE

You can choose one of the weighing units mentioned in the previous section as the power-on

weighing unit.  The selected unit will be the default unit each time when the scale is powered

on (other enabled units can be displayed in revolving order by pressing MODE.)

- Press MC/SET until display shows  Init = g.

- Press M+/MR to choose the unit you want to set as power-on operation unit.

DIGITAL FILTERING

This function is provided for weighing vibrating objects, or to prevent the scale from vibration

that may affect stability of the scale.

- Press MC/SET until display shows  FiL. X.				Fil. 1

- Press M+/MR to select the appropriate level for your application.  (Each level of setting is

a trade-off between reacting speed and the effect to filter vibration.  Larger number

means slower reacting speed and better effect to filter vibration, and vice versa.)

AUTO POWER-OFF

When auto power-off is enabled, scale will be automatically switched off after there has been

no load on the pan (and, scale is stable at zero) for the time duration set.

- Press MC/SET until display shows Aut.XX. (“XX” is the current setup of auto power off

time.)

- Press M+/MR to select auto power-off time as one of 5, 10, 30, 60 minutes or no (auto

power-off is disabled.)

BACKLIGHT

Backlight is provided for having better visibility when the scale is used in dark environment.

- Press MC/SET until display shows Iit.XXX.				Iit.Aut

- Press M+/MR to select backlight mode as Auto, OFF, or On.

Aut :  Backlight will be ON whenever there is load (greater than 9 x display resolution) on the pan and OFF when there is no load.

On :	Backlight will be always ON after scale is powered on.

OFF :	Backlight will be always OFF after scale is powered on.

ZERO BAND

Zero-band function is used to define the weight range within which the scale will remain zero

indication whenever there is disturbance.  Weight range is expressed in terms of number of

display divisions, both for positive and negative directions.

- Press MC/SET until display shows ZEro.X.				ZEro.1

- Press M+/MR to choose the range. 	Six options; 0d (d0), 1d (d1), 2d (d2), 3d (d3), 4d (d4), and 5d (d5) are available.

- Press ZERO to go to normal operation mode after all the setup is completed.

2-3 Offset value and keypad check

Offset value mode is provided for maintenance people to observe the internal offset value when it is needed.  In the offset value mode, to press respectively each key on the keypad can also perform check of the key’s function.  Follow the steps below to enter the offset value mode:

LAC Series

- Switch off the scale.

- Keep ZERO pressed and switch on the scale, display will show  SEt  SCALE  Func..

Then, to press ZERO again, offset value will be shown in WEIGHT and UNIT WEIGHT

displays, and KEy18 will be shown in TOTAL COUNT display.  With weighing pan on the

scale, the correct offset values should range from 9000 to 21000.  Once offset value is

out of this range, add a resistor (50K~2.2M Ohm) between load cell terminals “E+~S+” to

reduce offset value to the normal or, add a resistor (50K~2.2M Ohm) between load cell

terminals “E+~S-” to increase offset value to the normal.

- Press TARE key will clear the offset value shown in UNIT WEIGHT display, while offset

value shown in WEIGHT display remains.  In TOTAL COUNT display, after pressing

TARE key, KEy19 will be shown, if the function of TARE key is normal.

- To press respectively each key on the keypad can perform check of the key’s function.

The correct display in TOTAL COUNT should be the following if the key pressed is normal:

| Key pressed | Correct display | Key pressed | Correct display | Key pressed | Correct display |
| --- | --- | --- | --- | --- | --- |
| 1 | KEy 1 | 7 | KEy 7 | M+ | KEy 12 |
| 2 | KEy 2 | 8 | KEy 8 | UWS | KEy 13 |
| 3 | KEy 3 | 9 | KEy 9 | MC | KEy 14 |
| 4 | KEy 4 | 0 | KEy 0 | CE | KEy 15 |
| 5 | KEy 5 | CK | KEy 10 | MR | KEy 16 |
| 6 | KEy 6 | . (decimal pt.) | KEy 11 | SMPL | KEy 17 |

- After all required has been done, press ZERO to initiate power-on self-test, and the scale will be ready for use.

LAW Series

- Switch off the scale.

- Keep M+/MR pressed and switch on the scale, display will show  POS 2.  Then, in

sequence press successively M+/MR, ZERO, MODE, and M+/MR, offset value will be

shown in the display.  With weighing pan on the scale, the correct offset values should

range from 3000 to 20000.  Once offset value is out of this range, add a resistor

(50K~2.2M Ohm) between load cell terminals “E+~S+” to reduce offset value to the

normal or, add a resistor (50K~2.2M Ohm) between load cell terminals “E+~S-” to

increase offset value to the normal.

- To press respectively each key on the keypad can perform check of the key’s function.

The correct symbol displayed on the left-side of the LCD display should be the following if

the key pressed is normal:

| Key pressed | Correct display |
| --- | --- |
| M+/MR | LO |
| MC/SET | OK |
| MODE | HI |
| TARE | A blinking battery symbol; disappear right away |

- After all required has been done, press ZERO twice to enter normal weighing mode.

2-4 Power supply voltage

On the bottom of the scale, near the power core plug, there is a slide switch used to select

power supply voltage between 110V AC and 220V AC.  Be sure the power supply voltage is

set correctly.  Fuse in located also on the bottom, near the operator side.  Use slotted

screwdriver turning counterclockwise to take out the fuse set, and turning clockwise to plug

the fuse set in.

2-5 Recharging the battery

These scales can operate continuously by connecting directly to AC power or, instead of,

operate by the rechargeable battery.  When the battery voltage is low, a battery sign will start

blinking in each display window, power supply will be turned off automatically 8 hours after the

battery signs start blinking, for avoiding inaccurate measurement or calculation.

Do recharge the scale when the battery voltage is low, otherwise the battery may be damaged

due to over-discharged.  When the scale is being charged, the LED indicated by the word

CHARGE on the front panel will be in red, the LED will become in green when battery has

been charged to full capacity (charging time is approximately eight hours.)

2-6 RS-232 serial port option

Output data

| Baud rate | ：2400 , 4800 or 9600 bps |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Data bit | ：7 |  |  |  |  |  |  |  |  |  |  |  |
| Parity bit | ：E(even) |  |  |  |  |  |  |  |  |  |  |  |
| Stop bit | ：1 |  |  |  |  |  |  |  |  |  |  |  |
| Code | ：ASCll |  | LSB |  |  |  |  |  |  | MSB |  |  |
| Bit format | ： |  | 0 | 1 | 2 | 3 |  | 4 | 5 | 6 | 7 | 8 |

DATA BIT

START BIT                                   PARITY BIT  STOP BIT

Data format：

□□□□□□□□□□□□□□□□□□□□□□□□□□

±

weight data               unit weight			total counts    CR  LF

status		   ( 7 digits )			   ( 6 digits )			( 6 digits )

blank                  blank

| STATUS definition: (LAC) | ST US MR CK OL | = Stable = Unstable = Accumulation = Check counts = Overload |
| --- | --- | --- |
| STATUS definition: (LAW) Header 1: Header 2: | ST US NT GS | = Stable = Unstable = Net = Gross |

Example for output data：

ST+		120.00		0.200		600

UN+	      90.00	     0.300		300

Example of program in BASICA for executing data output：

10. OPEN ”COM1:9600,E,7,1,RS,DS,LF” AS#1

20. INPUT#1,A$, B$, C$, D$

30. PRINT A$, B$, C$, D$

40. GOTO 20

50. END

CHAPTER 3 MAINTENANCE

3-1 Replacing the main printed circuit board

Note! : Any intention to replace PCB and/or to repair components on the board must

proceed cautiously and carefully; otherwise further irreversible damage(s) may

be caused.

- Switch off the scale.

- Remove the weighing pan --- (1)

- Loosen and remove the two fixing hexagonal screws --- (5)

- Remove the pan support --- (2)

- Loosen and remove the four fixing screws --- (33)

- Move up carefully the upper cover.  Pay attention that several wires are connected to the main PCB mounted on the upper cover.  Place the upper cover upside down for better handling during working --- (3)

- Use soldering iron to remove the load cell wires from the board (E+, E-, S+, and S- marked on the PCB) --- (28)

- Remove the (J9) keyboard wire from the board --- (30)

- Use soldering iron to remove the green JGNP wire from the board --- (41)

- Remove the (J6) on/off switch wire connector from the board --- (24)

- Remove the (J4) battery wire connector from the board --- (26)

- Remove the transformer connector, (J3) for LAC, (J16) for LAW, from the board --- (32)

- Use soldering iron to remove the (J8) buzzer wires from the board (on the upper side of

the board; not shown in the figure)

- Loosen and remove two fixing screws --- (40)

- Take out the main PCB board --- (27)

- Prepare and install a new PCB

- Install the two fixing screws to the board --- (40)

- Solder buzzer wires back to J8 (red wire to + terminal)

- Solder the load cell wire (E+, E-, S+, S-) to the new board --- (27)

- Install the (J9) key board wires connector to the new board --- (30)

- Solder the green JGNP wire to the new board --- (41)

- Install the (J6) on/off switch wire connector to the new board --- (24)

- Install battery wire connector (J4) to the new board --- (26)

- Install the transformer connector, (J3) for LAC, (J16) for LAW, to the new board --- (32)

- Install the upper cover back --- (3)

- Install the four fixing screws --- (33)

- Install the pan support (directing a useless hole on it toward the level device) --- (2)

- Install the two fixing hexagonal screws --- (5)

- Install the weighing pan --- (1)

- Switch on the scale and check functions

3-2 Replacing load cell

- (Follow the steps 1 to steps 7 in the previous Section to remove load cell wires)

- Cut off the plastic wires-fastening-ring binding load cell wires and others

- Loosen and remove the two fixing hexagonal screws --- (36)

- Remove the load cell adaptor, pay attention not to alter the two load cell overload protection screws at the feet of the adaptor --- (11)

- Loosen and remove the two fixing hexagonal screws at the bottom --- (38)

- Take out the load cell --- (28)

- Keep the load cell spacer beneath the load cell at the original position --- (37)

- Prepare and install a new load cell on the spacer --- (28)

- Install the two fixing hexagonal screws at the bottom --- (38)

- Install the load cell adaptor --- (11)

- Install the two fixing hexagonal screws --- (36)

- Solder the load cell wires (E+, E-, S+, S-) to the board

- Use a new plastic wires-fastening-ring binding load cell wires and others for better clarity

- (Follow the step 25 to step 29 in the previous Section to complete)

- Switch on the scale and to check the offset value, and to conduct calibration to check.

CHAPTER 4 TROUBLESHOOTING

4-1 Preliminary checks

- Is the battery electricity capacity OK?

- Is AC power cord correctly plugged?

- Is the weighing pan the right one?

- Is there anything under the weighing pan?

- Is the scale at a flat and steady base?

- Are there any vibrating, rotating, and/or reciprocating equipment around?

- Is there any air turbulence within the environment?

- Is there any equipment that generates magnetic field around?

- Is there rapid temperature change within the environment?

- Are there any cooling and/or heating sources around?

4-2 Problems and Solutions

Note! : Any intention to replace PCB and/or to repair components on the board must proceed

cautiously and carefully; otherwise further irreversible damage(s) may be caused.

| Problems | Possible Causes | Solutions |
| --- | --- | --- |
| No any function pressing ON/OFF switch to turn on | - Malfunction in on/off switch - Malfunction in battery - Fuse burn down - Malfunction in U6 (7660), Q8 (A733), and/or Q9 (A733) - Malfunction in Y1 (ceramic resonator) - Malfunction in main PCB | - Check/replace on/off switch - Check/replace battery - Replace the fuse - Check/replace U6, Q8, and/or Q9 - Check/replace Y1 - Replace the main PCB |
| No any function pressing any key | - Malfunction in key board - Malfunction in key board connector (J9) - Malfunction in diode (D6~D10) – LAC only - Malfunction in main PCB | - Check/replace key board - Check/replace the key board connector - Check/replace diode (D6~D10) – LAC only - Check/replace the main PCB |
| Malfunction in battery recharge | - Malfunction in battery - Malfunction in main PCB | - Check/replace battery - Check/replace the main PCB |
| Error in LCD display | - Malfunction in LCD - Malfunction in U2 (74HC4066) – LAC only - Malfunction in main PCB | - Replace LCD panel - Replace U2 - LAC only - Replace the main PCB |
| No backlight in LCD | - Backlight function is set off - Malfunction in BD1 3.Malfunction in BL1, BL2, and/or BL3 – only one of the three in LAW - Malfunction in main PCB | - Set on the backlight function - Check/replace BD1 - Check/replace BL1, BL2, and/or BL3 – only one of the three in LAW - Check/replace the main PCB |
| Malfunction in auto power-off function | - Malfunction in D5 (1N4148), R22 (33K), R23 (42.4K), R25 (47K), and/or Q6 (C1815) - Malfunction in U1 (4069) - Malfunction in main PCB | 1. Check/replace D5, R22, R23, R25, and/or Q6 2. Check/replace U1 3. Check/replace the main PCB |
| No sounding, while all others normal | - Malfunction in buzzer - Malfunction in Q19 (A733) | - Replace buzzer - Check/replace Q19 |
| Unstable zero indication | - Malfunction in load cell - Malfunction in CP6, CP7, CP 8, CP 9, CP 10, CP11, CP12, C22, C23, and/or C25 - Malfunction in U7 (OP07), U8 (OP177), and/or U9 (OP07) - Malfunction in main PCB | - Check/replace load cell - Check/replace CP6, CP7, CP 8, CP 9, CP 10, CP11, CP12, C22, C23, and/or C25 - Check/replace U7, U8, and/or U9 - Check/replace the main PCB |
| Not able to weigh up to the full capacity | - Load cell delivery protector is locked - Malfunction in load cell - Mistake in calibration - Malfunction in U6 (7660), U7 (OP07), and/or C24 - Malfunction in main PCB | - Loosen and remove the protector (refer to Section 1-3) - Check/replace the load cell - Re-calibrate the scale - Check/Replace U6, U7, and/or C24 - Replace the main PCB |
| Error message “-Err” or “Err” shown | 1. Load cell delivery protector is locked 2. Malfunction in load cell 3. Malfunction in U6 (7660), U7 (OP07), U8 (OP177), U9 (OP07), and/or U10 (4052) 4. Offset value is too high (“Err”), or too low (“-Err”) | 1. Refer to Section 1-3 for releasing the protector 2. Check/replace load cell 3. Check/replace U6, U7, U8, U9, and/or U10 4. Refer to Section 2-3 for offset value adjustment |

CHAPTER 5 SPARE PARTS

5-1 LAC- 3015 / LAC-1260

| Item | Part no | Part Name/Description | Quantity |
| --- | --- | --- | --- |
| 1 | 52-0101-0001 | LAC-LCD | 1 |
| 2 | 61-0003-0101 | LAC transformer | 1 |
| 3 | 61-0202-0000 | LAC Recharge battery (4.5Ah), LONG | 1 |
| 4 | 80-0119-0001 | LAC (B) Assembled PCB V2.0 | 1 |
| 5 | 51-0015-0300 | Load Cell 1022-20K (LAC-1260) | 1 |
| 6 | 51-0015-0000 | Load Cell 1022-35K (LAC-3015) | 1 |

5-2 LAW- 3015 / LAW-1260

| Item | Part No | Part Name/Description | Quantity |
| --- | --- | --- | --- |
| 1 | 61-0003-0101 | LAWR transformer | 1 |
| 2 | 61-0202-0000 | Recharge battery (4.5Ah), LONG | 1 |
| 3 | 80-0117-0002 | LAWR (B) Assembled PCB V1.6 | 1 |
| 4 | 52-0101-0200 | LAW (LCD) | 1 |
| 5 | 51-0015-0300 | Load Cell 1022-20K (LAWR1260) | 1 |
| 6 | 51-0015-0000 | Load Cell 1022-35K (LAWR3015) | 1 |