Toledo

8510 code d’erreur

| Error Code | # Description | Suggested Corrective Action |
| --- | --- | --- |
| E1 | PROGRAM MEMORY FATAL ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. REPLACE DISPLAY PCB |
| E2 | INTERNAL RAM MEMORY ## FATAL ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. REPLACE DISPLAY PCB |
| # E3 | NOVRAM MEMORY ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. REPROGRAM, RECALIBRATE 3. REPLACE DISPLAY PCB |
| E8 | COMMUNICATION LOSS WITH DLC | 1. CHECK DC SUPPLY VOLTAGES 2. CHECK INTERCONNECT HARNESSES FOR LOAD CELL AND PCB'S 3. REPLACE LOAD CELL |
| E9 | DLC OUT OF RANGE, UNDER CAPACITY | 1. CHECK FOR MECHANICAL BIND 2. CHECK DC SUPPLY VOLTAGES 3. REPLACE LOAD CELL 4. REPLACE DISPLAY PCB |
| E10 | DLC MEMORY VERIFY FATAL ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. REPLACE LOAD CELL |
| E11 | DLC ROM ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. REPLACE LOAD CELL |
| E13 | DLC EEPROM ERROR | 1. CHECK DC SUPPLY VOLTAGES 2. RECALIBRATE SCALE 3. REPLACE LOAD CELL |
| E16 | MATH ERROR | PRESS CLEAR KEY TO CLEAR ERROR. CHECK PROGRAM AND RECALIBRATE |
| E32 | INSUFFICIENT TEST USED FOR CALIBRATION | RECALIBRATE WITH AT LEAST 50 OF PROGRAMMED SCALE CAPACITY |
| E34 | TEST WEIGHT OVER CAP | CHECK CAPACITY PROGRAMMING. USE TEST WEIGHT LESS THAN CAPACITY |
| E35 | CALIBRATION ERROR | RECALIBRATE. SHIFT ADJUST CAN'T COMPENSATE MORE THAN 100% |