Toledo

8522

# Code d’erreur

| Error | Error Description | Corrective measures |
| --- | --- | --- |
| E1 | Fatal ROM memory error | 1. Replace EPROM and Chip Carrier. 2. Replace Main PCB. |
| E2 | Fatal internal RAM error | Replace Main PCB. |
| E3 | Setup memory corrupt | 1.Enter setup and reprogram 2. Replace Main PCB. |
| E4 | Fatal external RAM memory error | Replace Main PCB |
| E6 | Analog verify failure | 1. Recalibrate scale. 2. Replace Analog Module. 3. Replace Main PCB. |
| E7 | DigiTOL load cell, format error | 1. Recalibrate scale. 2. Replace load cell. 3. Replace Main PCB. |
| E8 | DigiTOL load cell, no data | 1. Check load cell cable wiring. 2. Check load cell excitation voltage. 3. Check 8522 with DigiTOL load cell simulator. 4. Replace load cell. 5. Replace Main PCB. |
| E9 | Load cell out of range | 1. Check base for obstructions. 2. Verify load cell wiring and voltage. 3. Replace Analog Module. 4. Replace load cell |
| E10 | DigiTOL load cell, RAM error | 1. Recalibrate scale. 2. Replace load cell. 3. Replace Main PCB |
| E11 | DigiTOL load cell, ROM error | 1. Replace load cell. 2. Replace Main PCB |
| E16 | Math Overflow | 1. Recalibrate scale. 2. Replace Main PCB |
| E21 | Incorrect scale capacity | Press CLEAR key, then enter correct capacity. |
| E24 | Illegal high range division | Press CLEAR key, then enter correct division size. |

## Toledo 8522 Code d’erreur suite

| ### E25 | Illegal mid range division | Press CLEAR key, then enter correct division size. |
| --- | --- | --- |
| E26 | Illegal low range division | Press CLEAR key, then enter correct division size. |
| E27 | Illegal Overcapacity Entry | Press CLEAR key, enter legal overcapacity value. |
| E32 | Insufficient test weight or insufficient signal from load cell | 1.Press CLEAR key to clear. 2.Check base for obstructions 3. Verify scale capacity programming. 4. Verify load cell wiring. 5. Replace Analog Module. 6. Replace load cell. |
| E34 | Test weight to large | Press CLEAR key, test weight must be less than 105% of capacity. |
| E35 | Incorrect Test Weight Increment | Press CLEAR key, reenter test weight value with correct increment size (X 1, X2, or X51. |
| E36 | Build to small for load cell capacity | Press CLEAR key, recalibrate with larger capacity |
| E38 | Poor build for counting | 1.Press CLEAR key to accept current build 2.Or recalibrate with larger capacity. |
| E40 XX | Bad Record, ID=XX | Press CLEAR key, bad record is automatically erased. |
| EEE Or -EEE | Scale not zeroed | 1. Press ZERO key. 2. Check base for obstructions. 3. Recalibrate scale. 4. Verify load cell wiring. |