Toledo 8530

# Codes d’erreurs

| Error | Error Description | Corrective Measures |
| --- | --- | --- |
| E1 | Fatal Rom Error | 1. Replace EPROM and Carrier. 2. Replace Main PCB |
| E2 | Fatal Internal RAM Error | 1.Replace Main PCB |
| E3 | Corrupt SETUP Memory Error | 1: Enter setup and reprogram. 2. Replace Main PCB. |
| E4 | Fatal External RAM Error | 1. Replace Main PCB. |
| E5 | Display Verify Error | 1. Replace Main PCB. |
| E6 XX | Analog verify failure in Power Cell XX | 1. Recalibrate scale. 2. Check pit power supply voltages. 3. Check load cell cable XX. 4. Replace Power Cell XX. |
| E7 XX | Data format error, Power Cell XX | 1. Recalibrate scale. 2. Replace load cell. 3. Replace Main PCB |
| E8 XX | No communication with Power Cell XX | 1. Check pit power supply voltages. 2. Check COM A and COM B voltages 3. Check 8530 Power Supply voltages. 4. Check all load cell cables. 5. Check 8530 Main PCB with PowerCell Simulator. 6. Replace Main PCB. |
| E9 XX | Power Cell XX is out of range | 1.Check weighbridge for mechanical bind. 2. Check for load cell cable damage. 3. Check pit power supply voltages. 4. Replace Power Cell XX. |
| E10 XX | RAM error in Power Cell XX | 1. Recalibrate scale. 2. Replace load cell. 3. Replace Main PCB |
| E11 XX | ROM error in Power Cell XX | 1. Replace load cell. 2. Replace Main PCB. |
| E13 XX | NOVRAM error in Power Cell XX | 1. Check pit power supply voltages. 2. Check load cell cable XX. 3. Replace Power Cell XX. 4. Replace Main PCB. |
| E14 | Battery Backed Ram Memory Lost | 1. Replace Battery 2. Check 8530 Power Supply Voltages. 3. Replace Main PCB |

Toledo 8530

## Code d’erreur suite

| E16 | Math Overflow Error | 1. Enter setup and reprogram. 2. Replace Main PCB. |
| --- | --- | --- |
| E21 | Incorrect scale capacity | Press CLEAR key, then enter correct capacity |
| E24 | Illegal high range division | Press CLEAR key, then enter correct division size |
| E25 | Illegal mid range division | Press CLEAR key, then enter correct division size |
| E26 | Illegal low range division | Press CLEAR key, then enter correct division size |
| E27 | ### Illegal Overcapacity Entry | Press CLEAR key, enter legal overcapacity value |
| E32 | Insufficient Test Weight | 1. Verify correct capacity has been entered. 2. Use more test weight |
| E34 | Test weight larger than capacity | Press CLEAR key, test weight must be less than 105% of capacity. Verify correct capacity has been entered. |
| E35 | Incorrect Test Weight Increment | Press CLEAR key, test weight value entered must match the increment size entered in setup. |
| E36 | Insufficient Counts or Build to small for load cell | 1. Press CLEAR key to clear. 2 Check weighbridge for mechanical bin 3.Verify capacity and increment size has been entered 4.Check shift adjust value in setup step (97), if shift adjust values = 0 then step (18) must be performed. 5. Recalibrate with larger test weight 6. Check pit power supply voltages 7. Verify that raw count output of Power Cells increases as weight is applied to the weighbridge. 8.Replace main PCB |
| E37 | Calibration Checksum Error | 1. Recalibrate 8530. 2. Check 8530 Power Supply Voltages 3. Replace Main PCB. |
| E38 | Poor build for counting | 1. Press CLEAR key to accept current build 2. Or recalibrate with larger capacity |
| E40 XX | Bad Record at Memory Address XX | Press CLEAR key, bad record is automatically erased |

#### Toledo 8530 Code d’erreur suite

| E41 | Stored weight in wrong weight units. (Ib/kg | Operator is attempting to perform an operation with stored tare weights or accumulation and the weight units have changed from lb to kg or vice versa. Press the Ib / kg to return to original weight units and repeat the store operation. |
| --- | --- | --- |
| E42 | Can't recall tare in net mode | Operator is attempting to recall a stored tare weight with the 8530 in the net weight mode. Either select a permanent register with a stored tare weight of zero or clear the tare before attempting to recall a stored tare weight. |
| E43 | Invalid accumulation | Net weight accumulation is selected, step [29 1 ] and operator is attempting to accumulate in the gross weight mode. |
| E50 | Illegal Power Cell Command | 1. Verify AC power and ground. 2. Check 8530 Power Supply Voltages 3.replace Main PCB. |
| EEE or -EEE | Scale not zeroed | 1. Press ZERO key. 2. Check weighbridge for mechanical bind 3.Recalibrate scale. |