TOLEDO

8582 Code d’erreur

| # Error | # Description | # Corrective Measures |
| --- | --- | --- |
| I/0 1 ERROR | I/O 1 Communication Error | Replace Logic PCB |
| PRINT ERROR | Printer Communication Error | Replace Logic PCB |
| I/0 2 ERROR | I/O 2 Communication Error | 1. If no Serial I/0 2 PCB is present, disable I/O 2 in setup. 2. if no Bar Code PCB is present, disable Bar Code in setup. 3. Check cable to Serial I/O 2 PCB 4. Replace Serial I/0 2 PCB. |
| BAR CODE ERROR | Bar Code Communication Error | 1. If no Serial I/O 2 PCB is present, disable those selections in setup. 2. Check cable to Serial 1/0 2 PCB 3. Replace Serial I/0 2 PCB. |
| PROM CHECKSUM ERROR | PROM Failure | Replace Logic PCB. |
| PROCESSOR ERROR | Internal RAM Error | Replace Logic PCB. |
| RAM FATAL ERROR | Internal RAM Error | Replace Logic PCB. |
| NOV RAM FATAL-ERROR | Shadow RAM Error | Replace Logic PCB. |
| NO SCALE RESP | No response from any scale. | 1. Check internal harnesses to all scales and scale PCB's. 2. Replace internal load cell or Logic PCB. |
| NOV RAM CHECK ERROR | Novrarn Checksum Error | Enter setup/Reprogram/Exit Setup. |
| EXTRN RAM CHK ERR | External RAM Checksum Error | 1. Confirm battery voltage (see Section VIII-D-5.3. this manual). 2. Re-enter setup information. |
| RAM COLD START | RAM Checksum Error | Enter setup/Reprogram/Exit Setup. |
| SCL X FORMAT ERROR | Data Format Error | 1. Power down/wait 15 sec./power up 2. Recalibrate Scale. |
| SCL X COMMON ERR | No Data Received to the Scale PCB from the Analog or Digital Load Cell. | 1. Power down/wait 15 sec./power up. 2. Check cables to scale base X. 3. Replace load cell or Scale PCB. |
| SCL X POS OUT ERR | Scale Overcapacity, Scale and Logic PCB Z. are communicating. | 1. Remove weight from scale platform. 2. Check load cell connections. 3. Recalibrate scale X. 4. Replace load cell or Scale PCB. |
| SCL X NEG OUT ERR | Scale Under Zero | 1. Check load cell connections. 2. Recalibrate scale X. 3. Replace load cell or Scale PCB |
| SCL X DOWNLOAD ERR | No Scale Response | 1. Recalibrate scale X. 2. Replace load cell or Scale PCB. |
| SCL X MEMORY ERROR | Digital Load Cell Memory Failure | 1. Power down/wait 15 sec./power up. 2. Recalibrate scale X. 3. Replace load cell. |
| SCL X OVER ZERO | Scale is over zero and out of zero capture range. | if an error occurs just after calibration refer to section VIII-D-5.4. If this occurs during operation recalibrate the scale. |
| SCL X AV FAIL | Scale has failed the analog verify test three consecutive times. | 1. Cycle the 8582 AC power. 2. If the error re-occurs. enter set-up and recalibrate the scale. 3. Cycle power again. 4. It it fails AV test again check scale for linearity problems. |
| SCL X BOARD ERROR | The memory content has on the Scale PCs has failed it's checksum test | 1. Cycle power to the 8582. 2. If the error re-occurs replace the Scale PCB. |

TOLEDO

8282 Codes d’erreur suite

| SCALE X OUT | No Scale Response from Scale PCB to Logic PCB. | 1. Check load cell connection/Refer to Section VIII-D-6. 2. Power down/wait 15 sec./power up. 3. Replace load cell or Scale PCB |
| --- | --- | --- |
| XXX)(RECORDS BAD | Detective Files | This indicates file checksum failure. Bad records are automatically cleared by 8582. |
| FILE FULL | Memory Full | Indicates maximum number of records have been entered. |
| ACCUM FULL | Accumulator Overflow | Indicates the piece count accumulator has exceeded 8 digits |
| ACCUM UNDER | Accumulator Underflow | Indicates the accumulator has exceeded its capacity negatively |
| NOT SAVED | Transaction Not Saved | Indicates the transaction buffer is full. |
| NOT FOUND | Record Not Found | Inventory file with entered ID not found. |
| APW CHK FAIL | APW Tolerance Failure | Indicates APW tolerance check has failed. Press ENTER to accept or any other key to restart count |
| TRANS ACC OVR | Single Accumulator Overflow | Clear Accumulator. |
| CHECKING MEM | Memory Check | Displayed while memory checksums are checked at power up |
| LOW SAMPLE | Insufficient Sample Weight | Add additional samples. |
| COUNT OVER | Piece Count Over | Indicates piece count has exceeded the display capacity of 6 digits. |
| OVER ENH RNG | Over Enhancement Range | Indicates APW cannot be enhanced unless pieces are removed |
| T-LAN ERROR | Communication Error | Multidrop host interface has a transmission error. |
| RAM CHK ERROR | RAM memory on the Logic PCs is defective. | 1. Cycle power to the 8582. If error reoccurs replace logic PCB |
| OPT FILE MEM ERR | The memory content on the Memory-I/0 PCs has failed it's checksum test | 1. Cycle power to the 8582. 2. If no Optional Memory I/O PCB is present, disable Optional Memory I/0 PCB in setup. 3. If the error re-occurs replace the Memory-I/0 PCB. |
| XXXX MATH OVERFLOW | Typically a problem with calibration. | 1. Enter Setup/Reprogram/Recalibrate/Exit Setup 2. If error reoccurs. record the XXXX number and call an authorized Toledo Scale Service Representative. |
| LC TEST MODE | Power to the 8582 has cycled with the setup switch on. | Power to the 8582 has cycled with the setup switch on. |