##### METTLER TOLEDO

##### JAGUAR

##### CODE D’ERREUR

Lumière (Status Lights) sur board analog

| A | B | Status |
| --- | --- | --- |
| Blink | Blink | OK |
| On | On | RAM Error |
| On | Off | EPROM Error |
| Off | On | MELSI Error |
| Off | Blink | MMR Cell Error |

## Error Codes and Action

| Error Message | Description | Probable Cause | Remedy |
| --- | --- | --- | --- |
| AB_BAD_DISP_MODE | Bad Display Mode Command in Allen-Bradley Block Transfer Message | PLC Programming Error. | Correct PLC Program. |
| AB_BAD_FLD_NAME1 | Bad Shared Data Field Name in field 1 of Allen-Bradley Block Transfer Message | PLC Programming Error. | Correct PLC Program. |
| AB_BAD_FLD_NAME2 | Bad Shared Data Field Name in field 2 of Allen-Bradley Block Transfer Message | PLC Programming Error. | Correct PLC Program. |
| AB_BAD_FLD_NAME3 | Bad Shared Data Field Name in field 3 of Allen-Bradley Block Transfer Message | PLC Programming Error. | Correct PLC Program. |
| AB_BAD_FLD_NAME4 | Bad Shared Data Field Name in field 4 of Allen-Bradley Block Transfer Message | PLC Programming Error. | Correct PLC Program. |
| AB_WRITE_DISABLE | PLC attempt to write a shared Data Field that is write protected. When the JAGUAR is legal-for-trade mode, restricted Shared Data field cannot be written | PLC Programming | If necessary,you can access restricted parameters by removing the Legal for Trade jumper. |
| ALC_EE_CHKSM_ER | A Checksum Error was dectected in accessing the scale. Calibration parameters on the Analog Load Cell Card. | Static, power problems, inductive noise. Bad EEPROM. | Reset to Factory in Scale Interface Menu. Recalibrate scale. If problem persists, replace the Analog Load Cell card. |

| ALC_EE_NO_ACCESS | The JAGUAR terminal cannot access Scale Calibration parameters on the Analog Load Cell card | You have configured the JAGUAR terminal for a nonexistent Analog L/C card ; the Analog L/C card is not jumpered properly; the Analog L/C card is not saeted properly; the Analog L/C cardis not working | Check your configuration; Check the jumpers on the L/C card; Reseat the L/C card. If none of these actions correct the problem, put in a new Analog L/C card. |
| --- | --- | --- | --- |
| ALC_EEPROM_ERROR | Analog L/C EEPROM memory error. | Static, power problems, inductive noise. Bad EEPROM | Re-power and recalibrate. Check for good power, sup press noise; take static pre cautions. Replace Analog PCB |
| ALC_EPROM_ERROR | Analog L/C EPROM memory error | Defectivw Analog PCB | Replace appropriate Analog PCB |
| ALC_MELSI_ERROR | Analog load cell A/D error | A/D error has occurred. | Re-power the unit. Check with weight simulator. If error persists, replace Analog PCB |
| ALC_NO_RESPONSE | Analog load cell A/D communications error. | A/D error has occurred. | Re-power the unit. Check all ALC jumpers. Check with weight simulator. If error persists, replace Analog PCB |
| ALC_RAM_ERROR | Analog load cell A/D RAM error. | Static, power problems, inductive noise. Bad Analog PCB. | Re-power and recalibrate. Check for good power, sup press noise; take static pre cautions. Replace Analog PCB. |
| ALC_RESPONCE_ERR | Analog Load Cell A/D communications response error. | Internal error. | Re-power the unit. Chekc with weigh simulator. If error persists, replace Analog PCB |
| ALC_UNDEFINED_ERR | Analog Load Cell A/D undefined error | Analog Load Cell A/D memory error has occured | Verify progrmming and jumpers for Analog PCB are correct. If error persist in software prior to “B” revision upgrade software to “C” revision or later. |
| ARCNET_BAD_ADDRS ARCNET_DUP_ADDRS | You have configured the ARCNet address jumpers with either a duplicate address with another node on the network or an illegal ARCNet address. | The ARCNet address jumpers are not set up properly. | Check network address jumpers on main controller card. |
| ARCNET_TEST_ERR | ## The standard power up testing of the ARCNet adapter failed | ARCNet Adapter Failure. | Re-power up the JAGUAR terminal. If problem persists, replace the controller card. |
| BAD_NUMBER_CELLS | The JAGUAR terminal has been configured with an illegal number of load cells in a POWERCELL scale or DJ-Box scale | Improper Setup. | Check number of load cells configured for both scales. Corect the setup |
| BRAM Bad – Rst? | Battery backed RAM error | Setup parameters in Battery Back RAM have been corrupted. Most likely causes are too long of storage for the terminal, power removed from the terminal memory too long, battery failure, or hardware failure. | Respond (Y) yes to reset to factory settings. Reprogram setup parameters. If problem persists, you may have to replace battery, power supply, or controller card. |
| BRAM CKSUM ERROR | Setup variables corrupt. | Electrical malfunction.Power has been removed from the terminal memory too long. The battry and super-cap on the controller card has been drained. | Press ENTER to continue. Check setup parameters for desired setting. |
| BRAM Err – RST? Y (Will appear during power-up sequence only after loading new software.) BRAM VERSION ERR | Different program version detected. Storage locations for setup parameters and memory locations have been moved in a new software up date. | New software version has benn downloadded to the JAGUAR terminal. | Press ENTER to accept the reset default Y response. All parameters will be reset to factory default values. Reprogram the JAGUAR terminal setup parameters. |
| CALIBRATION_ERR | Calibration error. | Improper setup or calibration sequence, or bad load cell. | Check wiring. Check with simulator. Check load cell and recalibrate. Verify calibration setup parameters |
| Can't redim. var | JagBASIC programming error. | Once a JagBASIC application has declared a variable or an array, it can not later be redimensioned to a different size array. | Correct JagBASIC program |

| CHANGE PWCEL ERR | There was an error when attempting to change a POWERCELL address. | Communications error with POWERCELL. | Run the cell diagnostics to verify what the POWERCELL address. If it is still at its old address, try changing the address again. If the problem persists, you may have to replace the POWERCELL. |
| --- | --- | --- | --- |
| CLEAR-TARE-AT-0 | According to the scale setup parameters, the scale must be at gross zero in order to clear tare. | If you select Tare Inter- lock the scale must be at gross 0 in order to clear tare. | Check you local Legal For Trade requirements. If you don't want this feature, turn off the Tare Interlock selection. |
| Command error | An error occurred in trying to access a file from the JagBASIC interpreter. | Most likely, you tried access a file that does not exist. It is also possible that the file system has been corrupted. | Use the DIR command from the JagBASIC Interpreter to verify the directory of the RAM disk. If the file system has been corrupted, you need to re-initialize it from the JagBASIC setup menus and rebuild it from the backup files you are maintaining on a PC. |
| BRAM Power Fail! | The JAGUAR terminal detected low power supply voltage while attempting to write permanent data to BRAM Shared Data. | You have on early version of the JAGUAR terminal power supply or a bad power supply | Upgrade the power supply on the terminal with the latest version. |
| CONECT_NOT_ FOUND | Serial connection not found. | Improper serial setup. | Reset serial programming block to factory defaults. Re-program serial setup parameters. |
| CTL_EE CHKSM_ERR | Checksum error on accessing the EPROM on the controller Board. This EEPROM holds the calibration parameters for single-cell DigiTOL and DJBox scales. | Electrical noise, static discharge, or bad EEPROM chip. | Recalibrate DigiTOL scale. If problem persists, replace the EEPROM on the controller board or replace the controller board itself. |

| CTL_EE_NO ACCESS | Physical error on accessing the EEPROM on the Controller Board. This EEPROM holds the calibration parameters for cell DigiTOL and DJBox scales. | Hardware malfunction. | Power down/up the terminal. If problem persists, replace the EEPROM on the controller board or replace the controller board itself. |
| --- | --- | --- | --- |
| DEST_NOT_CONNECT | The remote JAGUAR terminal that is the destination for a cluster communications message is not connected. | The remote JAGUAR terminal that is the destination for a cluster communications message is not connected. | Verify network setup; verify ARCNet wiring, addresses, and terminations. |
| Device error | JagBASIC programming error. | The JagBASIC program has attempted to dimension a variable that is not an array. | Correct the Jagbasic program |
| Divide by zero | JagBASIC programming error. | The JagBASIC program has attempted to divide a number by zero. | Correct the JagBASIC program |
| DLC_ERR_NO_COMM | No communication or an intermittent failure to the DigiTOL load cell. | Bad DigiTOL base interconnect cable, or serial port. | Check voltages. Check with simulator. Verify serial output port. Check cable/cell. |
| DLC_BAD_PROTOCOL | The JAGUAR has detected a bad protocol exchange with a DigiTOL loadcell. | Noise being generated on cable between the terminal and DigiTOL load cell, | Check cabling, grounding, and connections at the terminal and at DigiTOL base. |
| DLC_INVALID_CHNL | DigiTOL load cell invalid channel. | Communications port as signed to the DigiTOL load cell is invalid. | Check COM port selection in setup. If setup appears correct, reset JAGUAR to factory setup. Reprogram the setup parameters |
| DLC_PARITY_ERROR | Parity error has been detected communication between JAGUAR and DigiTOL load cell. | Possible bad load cell, incorrect wiring, or electrical interference. | Check DigiTOL base, wiring, grounding, and power source. |
| EE A CErr - Rst? Y EE B CErr – Rst? Y | EEPROM Checksum Error. The scale calibration parameters stored on the EEPROM have been corrupted. | Hardware Failure. | Press ENTER to accept the reset default Y response. You must recalibrate JAGUAR scale. |
| EE A VErr - Rst? Y EE B VErr - Rst? Y EE VERSION ERROR | The version number in the EEPROM does not match the version expected by the JAGUAR operating system. | The JAGUAR terminal was calibrated with an earlier version of the JAGUAR operating system. | Press ENTER to accept the reset default Y response. You must recalibrate JAGUAR scale. |

| EE Reset Error | An attempt to access the EEPROM for the selected scale has failed. | Usually this is caused by an improperly configuring the JAGUAR with nonexistent scale devices. It can be also be caused by an improperly seated scale board or a hardware mat function on the scale board. | Check your scale configuration in setup. You may have to do Reset to Factory. Try reseating the boards. If these efforts fail, replace the scale board. |
| --- | --- | --- | --- |
| EF | Functional error in the weighing cell of an MMR (IDNET) base. | Static, power problems, inductive noise or unexpected operation. | Press the ESC key to continue. Take steps to eliminate probable cause. |
| EL | A command transmitted to an MMR (IDNET) base has been received, but cannot be executed. | Static, power problems, inductive noise or unexpected operation. | Press the ESC key to continue. Take steps to eliminate probable cause. |
| END_OF_FILE | End of File encountered while reading Shared Data. | End of File encountered while reading Shared Data. | None. |
| END_OF_SHIFT_ADJ | Last load cell or pair completed during shift adjust procedure | Last load cell or pair completed during shift adjust procedure | None. |
| ERROR in line | JagBASIC programming error. This message indicates the line in which the error occurred. | JagBASIC programming error. There will also be an error code indicating the type of programming error. | Correct the JagBASIC program. |
| ES | A command or string transmitted to an MMR (IDNET) base has been received, but is not a recognized command. | Static, power problems, inductive noise or unexpected operation. | Press the ESC key to continue. Take steps to eliminate probable cause. |
| ET | A transmission to an MMR(IDNET) base was received with a transmission error such as a parity, stop- or start-bit, or UART over flow error. | Static, power problems, inductive noise or unexpected operation. | Press the ESC key to continue. Take steps to eliminate probable cause. |
| Event def error | JagBASIC programming error. | There is a programming error in defining an event. | Correct JagBASIC program. |
| File open failed | JagBASIC programming error. | Most likely, the JagBASIC program has attempted to open a nonexistent RAM disk file or serial communications device. | Correct JagBASIC program. |

| FOS_RESP_ TIMEOUT | The Formatted Output Server(FOS) generates demand print and continuous print messages. They may be directed to a local or remote serial port. This error occurs when the FOS does not receive a response serial port driver within a specified amount of time. | This error usually occurs when print data is directed to a remote serial port. If the ARCNet LAN is disconnected while the FOS is waiting for a response, this error may occur. | Check ARCNet wiring. |
| --- | --- | --- | --- |
| IDN BUFF OVRFLOW | Excessive data, | Unexpected operation. | Power down, then up. If error still occurs, reset the JAGUAR terminal to Factory settings and repro gram setup parameters. |
| IDN EPROM ERROR | Damaged EPROM chip. | Electrical malfunction. | Replace IDNet PCB. |
| IDN NO RESPONSE | No response from base | Bad wiring or pcb(s). | Check wiring. Replace IDNet PCB. Replace IDNet base. |
| IDN RAM ERROR | IDN RAM ERROR | Electrical malfunction. | Replace IDNet PCB. |
| IDN RESPNSE ERR | Unexpected response. | Unexpected operation. | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters |
| Illegal command | JagBASIC programming error | The JagBASIC program has issued a command that is not a legal command. | Correct the JagBASIC program. |
| ILLEGL_ QUAR_RACK | An illegal quarter rack value has been specified for the Allen Bradley RIO option. | Invalid setup. | Check the Allen-Bradley set up |
| Incomplete line | JagBASIC programming error. | The JagBASIC contains a line that does not have the full syntax required for a line. | Correct JagBASIC program. |
| INCRM_CHAIN_TARE | A decreasing chain tare was attempted in a market where only incremental chain toring is permitted. | Chain taring that causes a decrease in the tare weight is not permitted in some markets in legal-for-trade applications. An incremental chain tare is a new tare on top of an already existing tare value where the new tare value is greater than the old tare value. | Check the market setting in setup. Check the "tare inter lock" setting in setup. Check the legal for-trade jumper on the controller board. Verify that these are set properly. |

| Internal Errors 1 Through 13 | Various errors. | Programming failure, hardware failure. | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. Replace controller or Analog PCB. |
| --- | --- | --- | --- |
| Invalid device # | JagBASIC programming error. | The JagBASIC program is referencing a device # that is not open. | Correct the JagBASIC program |
| Invalid SD name | JagBASIC programming error. | The JagBASIC program is referencing an invalid Shared Data name. | Correct the JagBASIC program |
| INVALID_FILE_NAME | There was an attempt to access Shared Data with an invalid file name. | This could be caused by an internal or external access of Shared Data | Try to determine if an internal or external access caused the error. If an external access is causing the error, then correct the PLC or Host PC program, If an internal source appears to be causing the problem, power down, then up. If error still occurs, reset JAGUAR to factory settings and reprogram setup parameters. |
| LADDER_EMPTY | Discrete I/0 setup error, | User attempted to delete a rung from the ladder that is empty. | Reset discrete configuration to factory and setup discretes again. |
| LADDER_FULL | Discrete I/0 setup error, | User attempted to add a rung to the ladder that is already full. | Reset discrete configuration to factory and setup discretes again |
| Line # invalid | JagBASIC programming error. | The JagBASIC contains a line number that is greater than 30000 or is a duplicate of an existing line number. | Correct the JagBASIC program. |
| Line too big | JagBASIC programming error. | The size of a JagBASIC line is greater than 80 characters. | Correct the JagBASIC program. |
| LOAD::no filename | JagBASIC programming error. | The LOAD command does not contain a file name. | Correct the JagBASIC program. |

| Memory find fail | JagBASIC programming error. |  | The JagBASIC program has exceeded the memory limits of the system. |  | There are many ways to reduce the memory usage of a JagBASIC program. Reduce number of lines. Eliminate unnecessary spaces in the program. Reduce the number of variables. Reduce the size of arrays. When chaining JagBASIC programs, you should always load in the largest program first. This will reduce memory fragmentation. |  |
| --- | --- | --- | --- | --- | --- | --- |
| NETWORK_XMIT_ERR | ARCNet communications transmission error. |  | Faulty ARCNet addresses wiring, line termination, or adapter |  | Check the ARCNet wiring for bad connections, wiring breaks, or improper line terminations |  |
| NEXT without FOR | JagBASIC programming error. |  | There is a NEXT statement without the required FOR statement. |  | Correct JagBASIC program. |  |
| No line number | JagBASIC programming error |  | The program line does not have a line number |  | Correct JagBASIC program. |  |
| No Remote Access | JagBASIC programming error |  | The program is attempting to access a device that is already in use by a serial connection or by another JagBASIC program in the JAGUAR terminal cluster |  | If you want to access a serial device, you must remove all serial connections to the de vice in setup. If you want to share a serial device among JagBASIC programs, you must set up a scheme where only one program has the de vice open at a time. |  |
| No Scale A Type No Scale B Type No Scale E Type | Scale type definition is missing |  | No scale type entered in Scale Interface menu |  | Go to the "Scale Interface" setup menu and properly set the scale type |  |
| NO_CHAIN_TARE | User attempted to take a second or "chain" tare after a tare was already taken. |  | When the tare interlock is selected in setup, chain taring is illegal in certain markets |  | Check the local "legal for trade" requirements. Check the market selection and tare interlock settings in setup. The system will continue to operate properly but will not allow the chain tare. |  |
| NO_DISCRETE_CBCK | There are no more discrete There are no more discrete |  | The setup of this system and the JagBASIC application has exceeded this system limit |  | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. |  |
| NO_DMD_PRNT_CON | There is demand print connection configured in setup |  | No demand print entered in the "Config Serial, Con figure Port" menu. |  | Change setup parameters |  |
| NO_KEYBOARD_TARE |  | Keyboard tare disabled. |  | Keyboard Tare is disabled in the "Application Envn,Tare Operation" setup menu. |  | Operation" setup menu. enable this feature. |
| NO_PUSHBUTT_TARE |  | Pushbutton tare disabled. |  | Pushbutton Tare is disabled in the 'Application Envn, Tare Operation' setup menu. |  | Change setup to enable this feature. |
| NO_PWCEL_OLD_ADR |  | POWERCELL readdressing |  | When readdressing a POWERCELL, no POWERCELL was found at the specified old address This problem could also be caused by a entering an invalid address or by a POWERCELL communications error. |  | Run the cell diagnostics to verify the POWERCELL ad dress. If the cell is at its old address, try changing the ad dress again. If the problem persists, you may have to re place the POWERCELL |
| NO_SECOND_UNITS |  | Secondary units not specified |  | No secondary unit selected in"Application Envn, Alt Weight Units" mode setup. |  | Change setup to enable feature. |
| OFF LINE ERROR |  | The JAGUAR terminal on ARCNET network not responding |  | Loss of communication between JAGUAR terminals |  | Check ARCNET cabling. Check JAGUAR terminal on network. Power down then up affected Jaguars. |
| ON no GOSUB |  | JagBASIC programming error. |  | ON statement is present without required GOSUB |  | Correct JagBasic program. |
| Oper. File Error |  | Cannot read language messages file. |  | Unexpected operation |  | Power down, then up. If error still occurs, reset JAGUAR to factory settings and reprogram setup parameters |
| OPTION BASE->DIM |  | JagBASIC programming error. |  | The JagBASIC program must define the OPTION BASE before dimensioning an array |  | Correct JagBasic program. |
| OUT_OF_COMM_BUFS |  | Cluster communications error |  | The system has exceeded the fixed limit on the number of communication buffers that can be used at one Most likely one JAGUAR is sending messages to a second JAGUAR faster than the second JAGUAR can process them |  | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. |

| OUT_OF_MEMORY | The JAGUAR software can not get the dynamic memory it needs to continue running | The system is using more dynamic "heap" memory than is available or the heap memory has become fragmented. | There are many ways to reduce the system memory usage, Reduce the size of a JagBASIC program. Eliminate unnecessary spaces in the program. Reduce the number of variables. Reduce the size of the arrays When chaining JagBASIC programs, you should always chain in the largest program first. This will reduce memory fragmentation Eliminate unused network connections, serial connections, and printer templates |
| --- | --- | --- | --- |
| OUT_OF_ZERO_RANG | Operator has attempted to zero the scale outside of the legal zeroing range | The zeroing limits are set up in the "Application Envn, Operation" menu | Change zeroing range in set up, if necessary |
| Overflow | JagBASIC programming error | A JagBASIC program causes an overflow error by exceeding certain system limits. In particular, the maximum size of the 'gosub' stack, the "for-next" stack, and the "while-wend" stack is9 entries each. Then, for example, if you try to subroutines more than 9 entries deep, you will get an Overflow error. Overflow errors can also be caused by particular language syntax errors | Correct JagBASIC program |
| PLC_COMM_ERROR | There was a timeout in the communications between the JAGUAR terminal and a PLC. This error is only reported after there has been a successful connection and the communication subsequently fails. | The most likely cause improper wiring or wiring termination. A fault at the PLC can also cause this error. Entering and exiting setup at the JAGUAR terminal can also cause this error because the PLC adapter is reset | Check wiring and wiring terminations. Verify PLC operation. Note: The JAGUAR terminal will automatically recover communications after an intermittent failure |
| Prbus config err | Attempts to configure the Profibus card failed. | Most likely, this error caused by a Profibus card hardware failure. | If problem persists, replace the Profibus card |
| PRINT REQUESTED | The operator has re- quested a Demand Print through the Control Panel. | None. | None. |

| PRINT_IN_PROGRES | The operator has requested a second demand print while the first is in progress. | None. | None. |
| --- | --- | --- | --- |
| PRINT_NOT_READY | Scale is in motion while attempting to print | None. | None. |
| Profbus init err | The Profibus card could not be initialized. | Bad Profibus card | Replace Profibus card |
| Program too big | JagBASIC programming error | First problem: The program exceeds 400 text lines or 18KB Second problem: This error can also occur while you are typing in a JagBASIC program when the temporary program buffer becomes full | For the first problem, separate the program into smaller files that can be run independently or chained together. When chaining, always start execution with the largest program to avoid memory fragmentation. For the second problem, save the current program and load it in again. This will cause a larger temporary program buffer to be allocated |
| PWC_PROTOCOL_ERR PWC_TIMEOUT_ERR PWC_UNDEFIND_ERR PWC_BUFF _OVFLOW | Communication Error between controller card and POWERCELL card. | Bad POWERCELL card | If error persists, replace the POWERCELL card. |
| PWC_CHECKSUM_ERR | Checksum error on firm ware on POWERCELL card. | Bad POWERCELL card | If error persists, replace the POWERCELL card. |
| PWC_EEPROM_ACCES | Unable to access EE PROM on POWERCELL card. | Bad POWERCELL card | If error persists, replace the POWERCELL card. |
| PWC_EEPROM_ERR PWC_EEPROM_CHECK | Checksum error on power scale calibration data stored on EEPROM on POWERCELL card | New version of JAGUAR software Hardware failure caused corruption of EEPROM data | Recalibrate scale. If problem persists, replace the POWERCELL card |
| PWC_NO_ERROR PWC_NO_ERROR2 | None. | None. | None. |
| PWC_RAM_ERR | RAM memory error POWERCELL card. | Bad POWERCELL | If error persists, replace the POWERCELL card. |
| PWCEL_AT_NEW_ADR | You attempted to readdress a POWERCELL to a new address that already exists on the POWERCELL | Your addressing procedure for the POWERCELL has created duplicate ad dresses. | Recheck your addressing scheme. If necessary, reset all POWERCELL addresses to the factory default address of 240. Begin addressing the cells again. Use the AutoAddress capability to minimize addressing errors. |

| PWCEL_BAD_FMT | The format of the data from the remote POWERCELL is invalid | Most likely, this is a communication error or power supply problem for the remote POWERCELL. Bad remote POWERCELL. | If problem persists, validate wiring, line terminations, and power in the POWERCELL network. Replace the remote POWERCELL, if necessary |
| --- | --- | --- | --- |
| PWCEL_EEP_ERR | The remote POWERCELL has reported a checksum error in its EEPROM. | Bad remote POWERCELL. | Replace remote POWERCELL |
| PWCEL_NEG RNG | The weight reported by a remote POWERCELL is in the negative weight range. | Bad remote POWERCELL. | Replace remote POWERCELL |
| PWCEL_NO_DATA | No weight data is being reported by a remote POWERCELL. | Most likely, this is a communication error or power supply problem for the re mote POWERCELLs. It could also be caused by a bad re mote POWERCELL | If problem persists, validate wiring, line terminations, and power in the POWERCELL network. Replace the remote POWERCELL, if necessary. |
| PWCEL_NO_RESP | The remote POWERCELL is not responding to polls from the JAGUAR terminal. | Most likely, this is a communication error or power supply problem for the remote POWERCELLs. It could also be caused by a bad re mote POWERCELL. | If problem persists, validate wiring, line terminations, and power in the POWERCELL network. Replace the remote POWERCELL, if necessary. |
| PWCEL ROM_ERR PWCEL RAM_ERR | The remote POWECELL is reporting an error in its local memory | bad remote POWERCELL. | Replace the remote POWERCELL |
| PWCEL_RESTART | The JAGUAR terminal has restarted a remote POWERCELL after the POWERCELL has not responded with valid data. | Most likely, this a communication error or power supply problem for the remote POWERCELL. It could also be caused by a bad remote POWERCELL. | If problem persists, validate wiring, line terminations, and power in the POWERCELL network. Replace the remote POWERCELL, if necessary. |
| Record not found | JagBASIC programming error | A record specified in GET statement for an sequential file could not be found in the file. | There should be an ON ERROR statement in the JagBASIC program to handle these potential situations |
| NETWORK PRINT ERR | There was a network error in attempting to print a demand print, continuous print, or setup report | This error typically when the demand print continuous print is directed to a remote JAGUAR. It occurs when ARCNet net work messaging fails. | Check network setup, addresses, wiring, terminations, and connections. |

| Resource in use | JagBASIC programming error. | The JagBASIC application tried to access a system resource that is already in use by another JAGUAR task In particular, a JagBASIC application cannot open a serial port that has been assigned to a serial port connection in setup. Also, when two or more JagBASIC applications are sharing a remote serial port, only one application can have the port open at a time. | Correct JagBASIC application. To share remote serial ports between multiple JagBASIC applications you will have to develop sharing logic that checks for this specific error code. |
| --- | --- | --- | --- |
| RETURN no GOSUB | JagBASIC programming error. | RETURN statement is present without required GOSUB. | Correct JagBASIC application. |
| SCALE_IN_MOTION | Scale in motion. This is a normal occurrence and not necessarily an error. | Motion on the scale during taring or zeroing the scale. | Try mechanical methods to stabilize the scale base first. Then, try changing the filtering to a stiffer setting in set up. Then, try changing the motion stability settings to make it less sensitive. |
| SCALE_UNDER_ZERO | The scale gross weight has gone more than "n" divisions below the current zero. The default "n" is 5, but it can be adjusted in set up. | The zero value for the scale could have been reset by hitting the zero but ton. There could be a connection problem to the base, particularly, with an analog base. | Take all weight off the scale base and reset the zero value. Zero settings in setup determine the range of how far from the calibrated zero that you can set a new zero value. If your weighing process uses below zero weight values, you can disable the under zero by setting the zero blanking value to 99 divisions. Check the analog base wiring. |
| SCL_OVERCAPACITY | The weight on the scale exceeds the calibrated capacity of the scale by more than 5 divisions. | There is too much weight on the scale based on calibration parameters. | Reduce the weight on the scale. |
| SD string > max. | JagBASIC programming error. | JagBASIC can only access Shared Data fields whose length is less than the maximum JagBASIC string size of 80 bytes. | Correct JagBASIC program. |

| SD_BAD_BUFFER | There was an invalid access to Shared Data. |  | A Shared Data access re quest provided a buffer that is not long enough. |  | Make sure that the external agency is providing a large enough buffer to match the requested Shared Data field. |  |
| --- | --- | --- | --- | --- | --- | --- |
| SD_WRITE_DISABLE | There was an invalid access to Shared Data. |  | An external agency at tempted to access a protected Shared Data field in legal-for trade mode. |  | Check external agency program. |  |
| SER CONST ERROR | The JAGUAR could not start or restart a logical serial connection. |  | JAGUAR software error. |  | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. |  |
| SER_BUFFER_FULL | The JAGUAR demand print buffer is full. |  | JAGUAR software error. |  | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. |  |
| SER_IN_TIMEOUT | There was a timeout waiting for serial input. |  | Most likely, the serial device talking to the JAGUAR terminal is has not sent the required input to the JAGUAR terminal. This could also be caused by communication errors. |  | Verify message exchange between the JAGUAR terminal and serial device, |  |
| SER_MSG_SEQ_ERR | There was an error in the sequencing of demand print messages. |  | Most likely, when one JAGUAR terminal is printing at a remote JAGUAR terminal, a message was lost in the ARCNet communications |  | If problem persists, check ARCNet wiring, terminations, and connections. |  |
| SERIAL_MSG_ERROR | The Serial Services software modules got an in valid request. |  | JAGUAR software error. |  | Power down, then up. If error still occurs, reset the JAGUAR terminal to factory settings and reprogram setup parameters. |  |
| SETPOINT_NO_RATE | The user has configured a rate setpoint, but has not configured rate calculation. |  | The user has not configured the rate calculation. |  | Configure the rate function in the Alternate Weight Units sub block. |  |
| SHIFT_ADJUST_ERRO | The shift adjustment factors could not be calculated. |  | The weight placements during the shift adjust procedure was incorrect |  | Carefully redo the shift adjustment, perhaps, with bigger weights. |  |
| SPX_NET_ERROR | SPX returned a network error status. |  | Most likely, there is a problem with the ARCNET wiring |  | Check the network setup, wiring, connections, addressing, and terminations. |  |
| Syntax error | JagBASIC programming error. |  | The JagBASIC program has a syntax error. |  | Correct the JagBASIC program. |  |
| TARE_ABOVE_LIMIT |  | Tare value exceeds the allowed limit. |  | In a legal for trade environment in certain markets, the tare value cannot exceed the highest weight in the lowest range of |  | Check local legal-for-trade requirements. |
| TARE_NOT_IN_INCR |  | Keyboard Tare is not entered in a rounded value to the nearest increment. |  | In certain markets, the keyboard tare value must be entered in as a value rounded to the nearest increment |  | Make sure the keyboard tare value is rounded to the nearest increment. |
| TARE_OVER_CAPCTY |  | Tare exceeds the capacity of the scale. |  | The tare value cannot exceed the capacity of the scale. |  | Make sure the tare value is less than the capacity of the scale. |
| TARE_TOO_SMALL |  | Pushbutton tare value is less than one division. |  | Weight on scale must be at least one division when taking Pushbutton tare. |  | Make sure scale has at least one division of weight before taking pushbutton tare. |
| TARE_UNDER_ZERO |  | Attempted to take tare when scale is under zero and has an invalid weight. |  | Cannot take tare when scale is under zero. |  | Make sure scale has valid weight before taking tare. |
| TEMPLATE_ERROR |  | Template error. |  | Error detected in template configuration. |  | Check template configuration. Correct it as necessary. If problem persists, reset template to factory and reenter template. |
| Too many dimens |  | JagBASIC programming error. |  | JagBASIC arrays can have at most three dimensions. |  | Correct the JagBASIC program. |
| T00_SMALL_INCRMT |  | Increment size is too small. |  | The scale increment size is too small so that you are asking for more resolution than the scale base is capable of supporting. |  | Choose a larger increment size parameter in setup and recalibrate the scale. |
| Type mismatch |  | JagBASIC programming error. |  | JagBASIC statement is using an invalid data type or is relating two in compatible data types. |  | Correct the JagBASIC program. |
| Undefined funct |  | JagBASIC programming error. |  | JagBASIC statement referring to on undefined function. |  | Correct the JagBASIC program. |
| Value out range |  | JagBASIC programming error. |  | The JagBASIC statement is referring to a value out of the range of acceptable values. |  | Correct the JagBASIC program. |
| WRONG_SCALE_MODE |  | Zeroing scale in net mode. |  | User attempted to zero the scale in net mode, |  | Clear tare to put scale in gross mode before zeroing scale. |
| ZERO_NOT_CAPTURED |  | Tare attempted before power up zero value was captured. |  | Tare attempted before power up zero value was captured. |  | Wait a few seconds after power up before attempting a tare. |