TOLEDO Mentor

| Pour entrer dans le menu : trouver le bouton de service sous le plateau ou à l'intérieur de l'appareil et peser dessus jusqu'à ce que SERVICE apparaissent à l'affichage suivit de TEST. se déplacer vers CALIBRAT à l'aide des touches et peser ensuite sur PRINT. suivre ensuite les étapes ci-dessous. |  |  |  |  |
| --- | --- | --- | --- | --- |
| STEP DESCRIPTION | PRESS KEY | DISPLAY |  | COMMENTS |
|  |  | LEFT RIGHT |  |  |
|  |  | CALIBR | LOCAL |  |
| 1. access the local scale calibration sub-block | PRINT | LOCAL | UNITS |  |
| 2. access the units options | PRINT | UNITS | LB (or currently select units) | the calibration unit refer only to the unit of the unit of weight you will use to calibrate the scale |
| 3. display the desired calibration units | TARE | UNIT | desired option | options include: lb, kg, gram. |
| 4. select the display calibration unit | PRINT | CAPACITY | 20 LB | the scale automatically proceed to the next parameter |
| 5. display the appropriate scale capacity | TARE | CAPACITY | desired capacity | options include: lb:10,15,20,30,37.5,60,120 gram:5000,6000,10000,12000,15000,30000,60000 kg: 5,6,10,12,15,30,60 |
| 6. select the displayed capacity option | PRINT | INCRMNT | 0,0002 or currently selected increment |  |
| 7. display the desired increment option | TARE | INCRMNT | desired increment | options are given based on the selected unit and capacity. |
| 8. select the displayed increment option | PRINT | EMPTY | SCALE | the scale now prompt you to the calibration process |
| 9. remove any residual weigh on the platter | PRINT | ADD | WEIGHT |  |
| 10. place the test weight on the platter | PRINT | DIGIT 1? | ? -------- | the scale is asking for the first most significant digit of the calibration weight to be entered. 5 digits must be entered. |
| 11.enter the first digit of the test weight value | TARE | DIGIT 1? | 0 -------- | TARE is used to scroll through all available numbers. Weight value entry begin at the left so 10 would be entered as 00010 |
| 12. when the correct value is displayed, press PRINT | PRINT | DIGIT 2? | 0? -------- | display moves to second digit to allow next numeric entry. |
| 13. enter the second digit of the test weight value | TARE | DIGIT 2? | 00------- | follow the same procedure for all 5 digits of test weigh entry. note that all digits must be entered |
| 14. enter the complete numeric value for the test weight on the platter | PRINT | BLANK | CAL DONE | calibration weights can now be removed from the platter |
| 15. continue to the OP Access sub-block. | PRINT | BLANK | OPACCESS |  |