Mettler Toledo

PantherPlus Code d’erreurs

| # Error | Description | Corrective measures |
| --- | --- | --- |
| E1 | PROGRAM MEMORY ERROR | Check power supply voltages. Replace main Logic PCB |
| # E2 | E2 INTERNAL RAM ERROR | Check power supply voltages. Replace Main Logic PCB. |
| E3 | EEPROM MEMORY ERROR | Check power supply voltages. # Reprogram. Recalibrate. Replace Main Lo is PCB. |
| E4 | EXTERNAL RAM ERROR | Replace Main Logic PCB. |
| E7 | A/D CIRCUIT MALFUNCTION OR NO ANALOG LOAD CELL CONNECTED | Program for correct load cell type. Check load cell and cables. Check power supply voltage Replace main logic PCB |
| E16 | # INTERNAL MATH ERROR | ## Press CLEAR to acknowledge Unit will reset |
| E20 | PREACT VALUE IS GREATER THAN SETPOINT VALUE | Clear preact value, then re-enter set point value |
| E32 | INSUFFICIENT TEST WEIGHT USED FOR CALIBRATION | Recalibrate using more test weight |
| E34 | TEST WEIGHT EXCEEDS | Use less than 105 % of capacity Press CLEAR and re_enter |
| E35 | SPAN CALIBRATION ERROR | Recalibrate if error persist ,check programming or replace load cell |
| E36 | # ANALOG LOAD CELL OUT OF RANGE | Recalibrate Replace load cell |
| E50 | WEIGHT CAN NOT BE DISPLAYED IN ALTERNATE UNITS | Some alternate units combinations are illegal. Choose another scale build or disable alternate units. |
| E60 | STACK OVERFLOW | Press CLEAR. Unit resets. |
| EEE | POSITIVE MORE THAN ZERO CAPTURE LIMIT OF 2%OF SCALE CAPACITY | Remove material from scale base. Disable AZM in setup. Cycle power |
| -EEE | NEGATIVE MORE THAN ZERO CAPTURE LIMIT OF 2%OF SCALE CAPACITY | Disable AZM in setup. Calibrate scale. Cycle power |