# History Records

At the completion of each Material Transfer operation, Q.iMPACT formats a History Record that contains the relevant details of the feed operation and attempts to send the record across a TCP/IP/Ethernet LAN connection.

The history record is ASCII data in the following format, where  (*hst* and *hst*) are starting and ending delimeters for the history record.

(*hst*^

channel record sequence number^

channel number^

material path^

start command^

start command status^

material transfer completion status^

completion status qualifier^

delivered weight^

deviation From target^

assembly realtime status 1^

assembly realtime status 2^

overlap group^

number of Secondary Overlapping feeds^

target weight^

positive tolerance^

negative tolerance^

material transfer ID & display message^

feed time^

weight units^

start time^

start weight^

end weight^

average flow^

average spill^

K1 constant^

K2 constant^

flow at cutoff^

actual spill^

consecutive good feeds^

total good feeds^

total bad feeds^

diagnostic status^

channel record sequence number^

entire Measuring Device block with ‘^’ separating each field^

entire Material Path block with ‘^’ separating each field^

*hst*)

- MEASURING DEVICE CHANNEL TABLE SHARED DATA

- Where the Hex number  ‘nn’ is the instance representing the Measuring Device 1-200.

Attr. 	Function/Value			Name	Format	Description

1	Measuring Device Assigned	unn01	UC	Setup. 1=Yes, 0=No

2	Q.iMPACT Terminal		unn02	C	Setup.  0, 1-9, A-K.  Cluster Terminal

containing Measuring Device.

3	Measuring Device		unn03	C	Setup.  Scale A – D or Flow Meter K – V.

4	Stable Scale Time		unn04	US	Setup. Number of seconds to wait for a

stable scale  reading before returning an

“Unstable Scale” failure status.  Range is 5-

10 seconds.

5	Reserved			unn05	US	Reserved.

- 6		Feed Override Time		unn06	US	Setup. Time in seconds that any external

- logic (such as Slow Stem Timer or an

- operator changing modes) will be inhibited

- from removing the permissive on the

- enabling logic that is controlling the Final

- Control Element (FCE). This is used to

- prevent  something other than the Fast-Cut-

- Off from closing the FCE, which might cause

- the PAC to develop erroneous data for

- subsequent update of its constants.

- The PAC sets a status indicating that the

- Material Transfer is within the feed override

- Time. Default = 20 seconds

- 7		Minimum Slow Step Time		unn07	US	Setup.  The PAC algorithm computes a “step

- slow” value.  If the computed value is less

- than the minimum specified here,  PAC uses

- the minimum instead. Typical range is 30-60

- seconds.

- 9		Overlap Feed Alone Time		unn09	US	Setup. Time that an overlapped feed must

- fed alone before cutoff, in seconds. Typically

- 10-20 seconds.

- 10		Overlap Time Tolerance		unn10	US	Setup. Additional time in seconds allowed for

- an overlapping feed to complete.

- 11		Zero Flow Threshold		unn11	Float	Setup. Flow rate below which the system

- assumes zero flow. Typically sized 0.001%-

- 0.01% of Max Scale Capacity/second.  In a

- “dump to empty”,  the PAC uses this value to

- determine when a “dump to empty” is

- complete.  After this point, PAC keeps the

- FCE energized until the “drain time“ expires.

- When this value < 0, PAC task uses absolute

- Value for drain timer for and aborts drain

- Timer at zero flow.

- 12		Unstable Device Flow Threshold	unn12	Float	Setup. Flow rate above which PAC

- generates a “Noisy Scale” condition when

- waiting for a stable scale reading. Typically

- sized 0.005%-0.05% of Max Scale Capacity

- per second. Must be larger than Minimum

- Flow.

- 13		Minimum Add			unn13	Float	Setup.  The smallest amount of material the

- system will attempt to add. Typically 0.01%-

- 1% of Max Scale Capacity.

- 14		Maximum Unit Size		unn14	Float	Setup. Over capacity point of a vessel, in

- weight units.

- 15		Dump Trip Point			unn15	Float	Setup. Level below which PAC considers

- Dump To Empty complete. Typically,

- 0.001%-0.01% of Max Unit Size. Must be

- less than smallest batch size for the Unit.

- 16		Min PAC Flow Rate Threshold	unn16	Float	Setup.  PAC starts to apply the predictive

- algorithm when measured flow exceeds this

- value.  Typically, 0.1% of Max Unit Size.

- Must be larger than Minimum Flow.  PAC

- sets the setpoint target = SP – Spill until the

- the flow rate reaches the this value.

- 19		Reserved			unn19	Float	Reserved.

- 30		Current Zero			unn30	Float	Calculated.  After a “Dump to Empty”

- operation, PAC sets this value to actual

- scale weight when it algorithmically detects

- the zero.  It helps track a “heel” buildup in

- the vessel.

MATERIAL-PATH TABLE SHARED DATA	(External Read/Write)

Where the ‘nn’ is the instance representing the internal Unit number 1-1000, in base 36 number representation.

Attr. 	Function/Value			Name	Format	Description

1	Material-Path Assigned		qnn01	UC	Setup. 1=Yes, 0=No

2	Instrument Channel		qnn02	UC	Setup. Channel 1 to 200 is the Q.iMPACT

logical measuring device channel in the

cluster that processes this Material-Path.

3	Text Descriptor			qnn03	19C	Text Descriptor for Material-Path ID

5	PAC Algorithm			qnn05	UC	Setup.  PAC Algorithm.

0 = Spill Only,

1 = K1,

2 = K2,

3 = Dump To Empty.

6	Feeder Type			qnn06	UC	Setup.   Feeder Type

- Gain In Weight,

- Loss In Weight,

- Flow Meter

- Hand Add

7	Slow Step Timer Factor		qnn07	F	Setup.  The Slow Step Timer is defined

Factor*( target / average flow ).  The factor

is normally set as 1.5, but is adjustable on a

material basis.

If SST Factor< 0, used absolute value for

SST Factor, but generate alarm only

When SST expires.

9	Minimum Open Time		qnn09	US	Setup.  Time in seconds that spill

compensation is not applied immediately

following control device opening. If = 0, no

minimum.   It is the minimum time in

seconds that a feed must be active

before it is considered “successful” and

will have its PAC parameters updated.

This check guarantees that the flow rate

is valid before PAC updates the PAC

parameters for a Material Path.

10	Drain Time			qnn10	US	Setup.  Time in seconds that the system will

wait for material to drain into a vessel after

feed is complete and before testing for

material delivery tolerance.

11	Average Flow Rate Low Limit	qnn11	Float	Setup. Lower alarm limit for the Average

Flow Rate “A”. Typically set at 50% of the

flow rate, in weight or volume units/second.

12	Average Flow Rate High Limit	qnn12	Float	Setup. Upper alarm limit for the Average

Flow Rate “A”. Typically set at 150% of the

flow rate, in weight or volume units/second.

13	Average Spill Low Limit		qnn13	Float	Setup.  Lower alarm limit for the Average

Spill “AA”. Typically set at 50% of the

average spill, in weight or volume units.

14	Average Spill High Limit		qnn14	Float	Setup.  Upper alarm limit for the Average

Spill “AA”. Typically set at 150% of the

average spill, in weight or volume units.

15	Algorithm Update Parameter	qnn15	Float	Setup. PAC uses this value in calculation of

Average Flow Rate “A”, Average Spill “AA”,

and Cutoff Constants “B”, “BB”, and “C”.

It controls how quickly the system

responds to a change in operating

conditions. Range is 0.0–1.0. Smaller values

(0.1–0.3) should be used for systems that

have few changes, while larger values (0.6-

0.8) are used for systems that change

frequently. Default value is 0.2.

16	Flow Rate Filter Sample Period	qnn16	UC	Setup.  This value determines specifies the

period of time, from 1 to 60 seconds, over

which the rate is calculated.   For lower

values,  PAC responds more quickly to

changes in rate.  For larger values, the rate

value changes more smoothly.

17	Destination Channel		qnn17	UC	Setup. Destination channel.  Range 0 –

200. If the Destination Unit = 0, PAC

disables Unit Verification and Batch Count

Checking. At beginning of the feed, the PAC

finds the weight in the destination channel to

determine if this Material Feed will cause an overflow in the destination channel.  At the end of the feedings, the PAC determines if the loss in weight in the source channel matches the gain in weight in the destination unit.  Note that there may be multiple concurrent feeds into the destination channel. Destination Channel = 255 indicates that the Destination is outside of this cluster.

18	Max Flow Rate Alarm Value.	qnn18	Float	Setup.  Flow rates above this value generate

and alarm and terminate the feed.  Value = 0

turns off alarm checking.

19	Average Flow Rate “A”		qnn19	Float	Calculated/ Initialized in Setup. Average

Flow Rate “A”, in weight or volume units /

second.

20	Average Spill “AA”		qnn20	Float	Calculated/ Initialized in Setup. Average Spill

“AA”, in weight or 	volume units.

21	Cutoff Constant K1		qnn21	Float	Calculated.  First-order cutoff parameter

used in spill calculation.

22	Cutoff Constant K2		qnn22	Float	Calculated. Second-order cutoff parameter

used in spill calculation.

23	Flow at Last Cutoff “Q”		qnn23	Float	Calculated. Material flow rate at the cutoff of

the last feed cycle, in weight or volume units

/ second.

24	Spill at Last Cutoff		qnn24	Float	Calculated. Material spill following the last

feed cycle, in weight or volume units.

25	Cutoff Constant B			qnn25	Float	Calculated.  PAC uses “B” in K2 model. It is

the “Average Flow Squared”.

26	Cutoff Constant BB		qnn26	Float	Calculated.  PAC uses “BB” in K2 model.  It

is the “Average Flow * Spill”.

27	Cutoff Constant C			qnn27	Float	Calculated.  PAC uses “C” in K2 model.  It is

the “Average Flow Cubed”.

28	Consecutive Good Feed Counter	qnn28	Long	Calculated. Counts consecutive good feeds.

29	Total Good Feed Counter		qnn29	Long	Calculated. Counts total good feeds.

30	Total Bad Feed Counter		qnn30	US	Calculated. Counts the total bad feeds.

Feeds aborted early in the Material Transfer

cycle are not counted as bad feeds.

31	Diagnostic Status Bits		qnn31	Long	Calculated. PAC sets feed diagnostic status

word at  completion of last feed.  A bit SET

means:

- OK to process requested feed

- Minimum Open Time passed OK

- Minimum Feed Time passed  OK

- Predictive algorithm started  OK

- Predictive algorithm completed OK

- Feed Complete & Waiting for Stable Scale

- Over Tolerance

- Under Tolerance

- Flow Rate within limits

- Flow rate within 50% of range

- Flow rate within 75% of range

- Spill within limits

- Spill within 50% of range

- Spill within 75% of range

- Primary Feed Alone

- Awaiting Start Stability

- Dump Trip Point Passed

- Validating Aggregate Feed

- Checking Fast Flow Rate Alarm

- Hand Add Complete Detected

- Feed Aborted Awaiting Drain

- Power Failure During Feed

- Configuration Reset During Feed

- Entered Setup During Feed

- ETC Error on Last Primary Feed

- – 31 Reserved

32	Last Material Transfer Status	qnn32	UC	Calculated. Last Material Transfer Status is

set at completion of feed.  Refer to arn06.

33	Cutoff Constant A_hat		qnn33	Float	Calculated.  PAC uses in K2 model.

34	Cutoff Constant AA_hat		qnn34	Float	Calculated.  PAC uses in K2 model.

### ####  END OF HISTORY RECORD  ####

Other Misc Info re Historical data

PAC Process History Data

PAC makes an entry in the Process History record at the completion of each Material Transfer for the PAC Process.   There are two data records for each PAC Process. These Shared Data fields allow the PAC Historian to retrieve the history record for each process.

ControlNet Class Code: 91 hex

PROCESS HISTORY SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	History Record 1			hst01	400C	Process History Record #1, Process #1.

Contains the following fields in ASCII format

Separated by ‘^’ characters in the following

order:

sequenceNumber, channelNumber,

materialPath, command, status,

lastMaterialTransferStatus,statusQualifier,

deliveredWeight,  spill_deviationFromTarget ,

realtimeStatus1, realtimeStatus2,

slowStepTimer, estimatedTimeComplete,

overlapGroup,

numSecondaryOverlappingFeeds,

targetWeight, positiveTolerance,

negativeTolerance, displayMessage,

feedTime, weightUnits,  startTime,

startWeight, endWeight .

### PAC Process Table

The Q.iMPACT runs the PAC Processes once a second.  When it runs, the PAC Process sets its current status in the Process Table Shared Data, as necessary.  In case of critical events, such as, the completion of a material movement operation or an error, the Q.iMPACT updates the status immediately.   The PAC Process also reports the first 8 fields in this table in the Q.iMPACT->Controller Cyclic Input Assembly.

ControlNet Class Code: 89 hex

An instance of the PAC Process Table Object may exist for each of 12 PAC Processes.  Each contains the numbered instance attributes.

- PAC PROCESS TABLE SHARED DATA

- Where the ‘n’ is the instance representing the internal PAC Process number 1-9, A, B, C.

Attr. 	Function/Value			Name	Format	Description

The Q.iMPACT sends the first 8 fields to Controller via cyclic communications.

- 1	Channel Number			ptn01	UC	Channel Number (1-200)

2	Real-time Status 1		ptn02	UC	Status 1 bit assignment.

- Communications Data Integrity Bit alternates Polarity Each Second

- Instrument Data Integrity OK

- Scale Over Capacity

- Scale Under Zero

- Scale Motion

- Material Transfer Cycle Active

- Discrete IO Final Control Element                0 = Off , 1 = On

- Waiting for Controller to Acknowledge Last Material Transfer/Hand Add complete

3	Real-time Status 2		ptn03	US	Status 2 bit assignment.

- Feed Type Bits1&2.

- Feed Type Bit  1&2.  0 = Gain In Weight, 1=Loss In Weight, 2 = Flow Meter, 3 = Hand Add.

- Manual (Not Auto) Mode

- Gross Weight Feed

- Feed Override Active – external logic inhibited from removing permissive

- Feed Failed

- Communication Error

- Device Stability Warning

- Very Unstable Device

- Too High or Too Low Flow at Cutoff

- Three Times Average Flow at Cutoff

- Fast Feed Rate Alarm

- Wait for All Overlap Requests

- Waiting to Start Primary Overlapped Feed

- Primary Overlapped Feed In Progress

- Secondary Overlapped Feed.

4	Feed Weight			ptn04	F	Feed Weight.  During feed, this field is the

Net Accumulated Weight.  At completion of

feed, this field is the Delivered Weight.  This

field is reset to zero at the beginning of a

feed.

5	Gross Weight			ptn05	F	Gross Weight.  For scales, this field the

gross weight.  For flow meters, this field is

the same as Feed Weight.

6	Rate				ptn06	F	Rate of change of weight in primary units per

second.

7	Slow Step Timer			ptn07	US	“Time until Slow Step Timer Expires” in

Seconds = Slow Step Factor * (target /

average flow)

8	Estimated Time to Complete	ptn08	US	“Estimated Time to Complete” in Seconds

The Q.iMPACT does NOT send the following fields to Controller via cyclic communications.

- 9	Current Material Path Index	ptn09	US	Current Material Path Index.

- 10	Current Command Number	ptn10	UC	Current Command Being Processed.

- 11	Current Overlap Group		ptn11	UC	Current Overlap Feed Group.

- 12	Current # Overlapping Feeds	ptn12	US	Current # Overlapping Feeds.

13	Current Target Weight		ptn13	F	Current Target Feed Weight.

14	Current Positive Tolerance		ptn14	F	Current Positive Feed Tolerance.

15	Current Negative Tolerance	ptn15	F	Current Negative Feed Tolerance.

16	Batch ID/ Display Message	ptn16	40C	Batch ID from Host Controller that is used for

Data Collection Messages.   If there is a “^”

in the field, data following the “^” is used

as a Display Message for the Q.iMPACT

display.

17	Feed Time			ptn17	Long	Feed time in seconds. PAC sets this current

value once a second during feed and sets

total time at completion of feed.

18	Weight Units			ptn18	3C	ASCII descriptor for weight units. “lb”, “kg”,

”g”,  or “t”

19	Diagnostic Status			ptn19	Long	PAC sets feed diagnostic status word during

feed cycle.  At completion of feed, the limits

values reflect status of feed.  Refer to qnn31

for bit meaning assignments.

20	Feed Start Time			ptn20	L	Time in seconds from 1970 GMT

21	Unit Weight at Start		ptn21	D	Unit Weight at Start of Feed

22	Unit Weight at Completion		ptn22	D	Unit Weight at Completion of Feed