1	SHARED DATA	4

2	SHARED DATA HEAP ELEMENTS	4

2.1	Scale Weight Object	4

2.2	Dynamic Data Object	5

2.3	JagBasic Application Interface to Remote Batching	8

3	SHARED DATA STATIC BRAM ELEMENTS	8

3.1	Scale Weight Static Object	8

3.2	Scale Calibration Object	9

3.3	Scale Tare Object	11

3.4	Setpoint Object	12

3.5	System Object	13

3.6	User Literals Object	15

3.7	User Prompts Object	16

3.8	User Variables Object	16

3.9	Cluster Variables Object	16

3.10	Print Templates	17

3.11	Security	17

3.12	Serial Ports	17

3.13	Network	18

3.14	Email Alert	19

3.15	FTP Login	19

3.16	Network Remote Nodes	19

3.16.1	Remote Node Addresses	19

3.16.2	Remote Node Enable	20

3.17	PLC Configuration	20

3.17.1	PLC Configuration Shared Data	20

3.17.2	PLC (Non-PAC) Scale Configuration	20

3.17.3	Analog Output	20

3.18	CONNECTION SHARED DATA	21

3.19	LADDER LOGIC DATA	21

3.20	EEPROM SAVE AREA	21

3.21	BASIC Application Object	21

3.22	Power Cell Log Object	23

4	SHARED DATA EEPROM ELEMENTS	23

4.1	Scale Calibration - EEPROM Object	23

4.2	EEPROM Version Identification	24

4.3	Shift Adjust Object	24

4.4	Cell Calibration Object	24

4.5	Scale Zero	25

5	GLOBAL DISCRETE I/O DATA	25

5.1	Level Sensitive Discrete Status Object	25

5.1.1	Status of Scale A	25

5.1.2	Status of Scale B	26

5.1.3	Status of Scale C	26

5.1.4	Status of Scale D	26

5.1.5	Status of Scale E (Summing Scale)	26

5.1.6	Status of Setpoints 1-12	27

5.1.7	Status of Remote Node Connections	27

5.1.8	JagBasic Applications Real-Time Status for a PLC or PC.	27

5.1.9	Command Status Bits for Scale A	28

5.1.10	Command Status Bits for Scale B	28

5.1.11	Command Status Bits for the Selected Scale	28

5.1.12	Command Status Bits for Custom Print	29

5.1.13	Command Status Bits for Scale C	29

5.1.14	Command Status Bits for Scale D	29

5.1.15	Command Status Bits for Scale E	30

5.1.16	Miscellaneous JagX O/S Triggers	30

5.2	Edge Sensitive Discrete Status Object	30

5.2.1	Miscellaneous JagX O/S Internal Triggers	30

5.2.2	JagX O/S Event Indications	31

5.2.3	Command Triggers for Scale A	31

5.2.4	Command Triggers for Scale B	32

5.2.5	Command Triggers for the Selected Scale	32

5.2.6	Command Triggers for the Scale C	32

5.2.7	Command Triggers for the Scale D	32

5.2.8	Command Triggers for the Scale E  (Summing Scale)	33

5.2.9	Select Scale Command Triggers	33

5.2.10	Custom Print Command Triggers	33

5.2.11	Miscellaneous Command Triggers	34

5.2.12	PC Host Real-Time Commands for a JagBasic Application	34

5.2.13	Reserved  Triggers for JagX O/S	34

5.3	Rising/Falling Input Edge-Sensitive Object	34

5.4	Level Sensitive Physical Discrete I/O Object	35

6	SPECIAL REMOTE ACCESS FILE NAMES	36

7	Q.iMPACT Shared Data	36

7.1	Global Cluster Configuration	36

7.1.1	Q.iMPACT Cluster Terminals Table	36

7.1.2	Q.iMPACT Terminal Enable Table	37

7.1.3	Q.iMPACT Terminal Online Table	37

7.1.4	Q.iMPACT Instrument Channel Configuration	38

7.1.5	Q.iMPACT Global Parameters Table	38

7.2	Q.iMPACT – Controller Communications HEAP Fields	39

7.2.1	PAC Communications Command Data	40

7.2.2	PAC Communications Command Status	40

7.2.3	Q.iMPACT->PLC Cyclic Communications Input Assembly	41

7.3	PAC Processes	42

7.3.1	PAC Process Command Table	43

7.3.2	PAC Process Command Status Table	44

7.3.3	PAC Process Table	46

7.3.4	PAC Process History Data	48

7.4	Measuring Device Channel  Table	49

7.5	Measuring Device Channel Defaults Table	50

7.6	Material-Path (i.e. Raw Material) Table	51

7.7	Material-Path Defaults Table	54

7.8	Flow Meter	54

7.8.1	Flow Meter Heap Weight Object	54

7.8.2	Flow Meter Calibration Object	55

8	CHANGE HISTORY	57

# SHARED DATA

Shared Data is the global data storage for the JagXtreme.  Shared Data “blocks” contain related “fields” of data. The Scale threads, the Control Panel Setup, and the Web Pages set many of the Shared Data fields.  Applications, such as JagBasic programs, the PC OPC Server, the PC Windows API, Save/Load Utilities, Allen-Bradley PLC’s, Modbus Plus PLC’s, and PROFIBUS PLC’s, may also read or write Shared Data.

When the JagXtreme detects that the Legal-for-Trade switch is ON, it prohibits applications from writing to the blocks marked below as “External Read Only”. When the JagXtreme detects that Legal-for-Trade switch is OFF, it does not enforce any application write access restrictions.

- ControlNet PLC’s can access the numbered instance attributes.

# SHARED DATA HEAP ELEMENTS

## Scale Weight Object

ControlNet Class Code: 64 hex

An instance of the Scale Weight Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SCALE WEIGHT SHARED DATA	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Displayed Gross Weight		wtn01	13C	12 alphanumeric, right-justified + 0

2	Displayed Net Weight		wtn02	13C	12 alphanumeric, right-justified + 0

3	Displayed Weight Units		wtn03	3C	3 alphanumeric (lb pounds, kg kilograms, grams, t metric tons)

4	Displayed Aux Gross Weight	wtn04	13C	12 alphanumeric, right-justified + 0

5	Displayed Aux Net Weight		wtn05	13C	12 alphanumeric, right-justified + 0

6	Displayed Aux Weight Unit		wtn06	6C	6 alphanumeric (lb pounds, kg kilograms,

oz ounces, lb-oz pounds & ounces, oztroy ounces, dwt penny weights, metric tons, ton, or custom units name)

7	Displayed Aux Rate Period		wtn07	C	1 alphanumeric  (No, Sec, Min, Hour)

8	Displayed Rate			wtn08	13C	12 alphanumeric, right-justified + 0

9	Displayed Diagnostic Weight	wtn09	13C	12 alphanumeric, right-justified + 0

10	Legal Gross Weight		wtn10	D	double float weight

11	Legal Net Weight			wtn11	D	double float weight

12	Auxiliary Gross Weight		wtn12	D	double float weight

13	Auxiliary Net Weight		wtn13	D	double float weight

14	Auxiliary Rate			wtn14	D	double float weight

15	Scale State			wtn15	UC	0 = disabled, 1 = normal weight processing,

2 = diagnostic, 3 = calibration,

4 = shift adjust, 5 = error.

16	Continuous Output Status A	wtn16	UC	1 byte - any value.

17	Fine Gross Weight		wtn17	D	double float weight

18	Fine Net Weight			wtn18	D	double float weight

19	Weighing Range			wtn19	UC	0=single range weighing

1=multi-range 1, 2=range 2, 3=range 3

20	WIM Time Counts		wtn20	D	double float WIM time counts

21	WIM Weight Units		wtn21	3C	3 alphanumeric (lb pounds, kg kilograms,

grams, t metric tons)

22	New Conveyor Item		wtn22	US	Flag that scale thread sets to alert JagBasic

application that a new item is entering

conveyor scale.

23	Conveyor Item Count		wtn23	US	Number of items on conveyor scale.

24	Conveyor Weighment Count	wtn24	US	Number of A/D updates in conveyor

weighment data.

## Dynamic Data Object

ControlNet Class Code: 65 hex

- The Dynamic Data Object has only a single instance, instance 1.

DYNAMIC DATA SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

- 1	Read the JagXtreme Display	bd001	50C	Read contents of JagXtreme display. The first 8 bytes have the contents of the numeric display.  The next 16 bytes have the contents of the alphanumeric display.  The next 8 bytes have the contents of  the numeric cursors.  The next 16 bytes have the contents of the alphanumeric cursors.  For each cursor, its corresponding byte has an ASCII ‘0’ if the cursor is off and an ASCII ‘1’ if the cursor is on.

Board Configuration String		bd002	60C	Contains a 15-byte entry for each of four board slots.  Each entry contains a two-character board identifier and a 13-character board software serial number where applicable.

3	LatestKeyStrokeKeySource	bd003	2C	2  Alphanumeric.  Read/Write.

EEPROMAuthorization String	bd004	10C	10 Alphanumeric.  Read Only.

ConsoleSoftwarePartNo		bd005	13C	12 Alphanumeric+0.  Read Only.

Analog 1 Software Part No		bd006	13C	12 Alphanumeric+0.  Read Only.

Analog 2 Software Part No		bd007	13C 	12 Alphanumeric+0.  Read Only.

DigiTOL 1 Software Part No	bd008	13C 	12 Alphanumeric+0.  Read Only.

DigiTOL 2 Software Part No	bd009	13C 	12 Alphanumeric+0.  Read Only.

Idnet 1 SoftwarePartNo		bd010	13C 	12 Alphanumeric+0.  Read Only.

Ident 2 SoftwarePartNo		bd011	13C	12 Alphanumeric+0.  Read Only.

PowerCell SoftwarePartNo		bd012	13C 	12 Alphanumeric+0.  Read Only.

JagBasicDisplay			bd013	24C	Reserved for JagX O/S Use Only.

HMI Weight Stream		bd014	105C	Contains HMI Weight streams for up to five scales.	The HMI subscribes to which fields the JagXtreme will send by sending the subscribe message.  Its format is: <STX>S<ABCDESJL><ETX><chk>, where

ABCDE represents the scales, S represents the selected scale, L represents the lower JagX Display, J is the JagBasic Message Field, and  T is Time.   “S” is mutually exclusive from ABCDE.   Jag60 stores subscription.  The HMI Weight Stream is formatted as follows: stream1><US><stream2><US><streamn>

Each Weight Stream has the following contents:

<JagX ID>   1N     Range: 1 to 6

<Scale ID>   1A     Range: A to E.

If selected scale, range is

in lower case <a to e>.

<Status>       1C

Bit 7        Always 0

Bit 6        Always 1

Bit 5        1 = Scale In Motion

Bit 4	 1 = Center of Zero

Bit 3-2      00 = single range

01 = weight range 1

02 = weight range 2

03 = weight range 3

Bit 1        1 = Net Mode

Bit 0	 1 = Preset Tare

<Units>   1N          0=None, 1=lb, 2=kg, 3=g,

4=oz, 5=ozt, 6=dwt, 7=t,

8=ton, 9=custom

<Net Wt>      8N    6 digits plus possible “-“

and “.”

<Tare Wt>    8N    6 digits plus possible “-“

and “.”

HMI JagBasic Subscription String	bd015	24C	Contains JagBasic Subscription String

for HMI.

16	Power Cell Scale-Cell Errors	bd016	25C	25 Bytes. There is an error number for up to 24 power cells. This field has cell errors for both Scale A, Scale B, Scale C, and Scale D.   Read Only.

Scan Table			bd018	25C	Scan Table contains ordered list of

current power cell addresses.  Read Only.

19	Power Cell Scale-Cell Counts	bd019	24L	96 Bytes.  Each long contains the current

shift-adjusted counts for consecutive power cells in a scale.  An application can request the current counts for a scale by setting trigger t_69d for Scale A, t_6ad for Scale B to 1, t_62d for Scale C, and t_63d for Scale D.  Read Only.

20	Power Cell Overload State		bd020	25C	25 Bytes.  There is one entry each for up to 24 power cells. 0 = Cell not assigned, 1 = Cell OK, 2 = Cell in Overload condition.

21	Power Cell Zero Drift State		bd021	25C	25 Bytes.  There is one entry each for up to 24 power cells. 0 = Cell not assigned, 1 = Cell OK, 2 = Cell in Zero-Drift-Threshold-Exceeded condition.

CntrlNet Software Part Number	bd022	13C	12 Alphanumeric+0.

Conveyer Scale Command		bd023	UC	JagBasic Application sets this flag to

send a command to the Analog Board

Conveyor Scale.  Reset, Begin, End Weigh.

JagBasic Email Destination	bd024	40C	JagBasic Email command loads this variable with the Destination string. (future)

JagBasic Email Subject Line	bd025	40C	JagBasic Email command loads this variable with the Subject line string. (future)

JagBasic Email Active Status	bd026	UC	JagBasic Email command sets status to 1 if email can be sent, else 0. (future)

30	Read Discrete Inputs		bd030	US 	Read all discrete inputs with one access.  The returned data field has one bit representing each discrete input.  Remote read only.

31	Read Discrete Outputs		bd031	US	Read all discrete outputs with one access.  The returned data field has one bit representing each discrete output.  Remote read only.

32	Read Status Flags for Scale A	bd032	US	Read scale status flags for Scale A remotely.

33	Read Status Flags for Scale B	bd033	US	Read scale status flags for Scale B remotely.

34	Read Status Flags for Scale C	bd034	US	Read scale status flags for Scale D remotely.

35	Read Status Flags for Scale D	bd035	US	Read scale status flags for Scale D remotely.

36	Read Status Flags for Scale E	bd036	US	Read scale status flags for Scale E remotely.

37	Read JagBasic Custom Flags	bd037	US	Read flags s_250 - s_s5f remotely.

38	Read Cal Switch			bd038	US	Read Cal Switch remotely.

Jag39 Alias			bd039	UC	Web Server alias to read Jag39 value.

Jag40 Alias			bd040	UC	Web Server alias to read Jag40 value.

Jag41 Alias			bd041	UC	Web Server alias to read Jag41 value.

Jag42 Alias			bd042	UC	Web Server alias to read Jag42 value.

Jag43 Alias			bd043	UC	Web Server alias to read Jag43 value.

Jag44 Alias			bd044	UC	Web Server alias to read Jag44 value.

Jag45 Alias			bd045	UC	Web Server alias to read Jag45 value.

Jag46 Alias			bd046	UC	Web Server alias to read Jag46 value.

Jag47 Alias			bd047	UC	Web Server alias to read Jag47 value.

Jag48 Alias			bd048	UC	Web Server alias to read Jag48 value.

Jag49 Alias			bd049	UC	Web Server alias to read Jag49 value.

Jag50 Alias			bd050	UC	Web Server alias to read Jag50 value.

Jag51 Alias			bd051	UC	Web Server alias to read Jag51 value.

Jag52 Alias			bd052	UC	Web Server alias to read Jag52 value.

Jag53 Alias			bd053	UC	Web Server alias to read Jag53 value.

Jag66 Alias			bd054	UC	Web Server alias to read Jag66 value.

Jag67 Alias			bd055	UC	Web Server alias to read Jag67 value.

HMI Attached			bd056	UC	The HMI must set this to 1 after communications are established.  The Q.iMPACT sets this to 0 at power-up. (future)

Unacknowledged Error Present	bd057	UC	The Q.iMPACT sets this to 1 when an error is displayed.  Acknowledging the error will set it back to 0. (future)

Flow Meter 1 Software Part No	bd060	13C	12 Alphanumeric+0.  Read Only.

Flow Meter 2 Software Part No	bd061	13C	12 Alphanumeric+0.  Read Only.

Flow Meter 3 Software Part No	bd062	13C	12 Alphanumeric+0.  Read Only.

ControlNet Ping Status		bd070	US	Status Values:

- PING_INVALID

- PING_POWER_UP

- PING_CHECK_FOR_CABLE

- PING_WAITING_TO_ROGUE

- PING_CHECK_MODERATOR

- PING_IM_ALIVE

- PING_ATTACHED_ONLINE

- PING_FORCED_LISTEN_ONLY

- PING_DUPLICATE_LISTEN_ONLY

- PING_CNET_TIMEOUT

DisplayBoard			bd085	US bit	1 = Yes, 0 = No.  Read Only.

Analog Scale Board 1		bd086	US bit	1 = Yes, 0 = No.  Read Only.

Analog Scale Board 2		bd087	US bit	1 = Yes, 0 = No.  Read Only.

Allen-Bradley RIO		bd088	US bit	1 = Yes, 0 = No.  Read Only.

PROFIBUS			bd089	US bit	1 = Yes, 0 = No.  Read Only.

ControlNet			bd090	US bit	1 = Yes, 0 = No.  Read Only.

Multi-Function IO 1		bd091	US bit	1 = Yes, 0 = No.  Read Only.

Power Cell Scale Board		bd092	US bit	1 = Yes, 0 = No.  Read Only.

ModBus Plus			bd093	US bit	1 = Yes, 0 = No.  Read Only.

Analog Output Board		bd094	US bit	1 = Yes, 0 = No.  Read Only.

High Precision Scale 1		bd095	US bit	1 = Yes, 0 = No.  Read Only.

High Precision Scale 2		bd096	US bit	1 = Yes, 0 = No.  Read Only.

Flow Meter Board 1		bd097	US bit	1 = Yes, 0 = No.  Read Only.

Flow Meter Board 2		bd098	US bit	1 = Yes, 0 = No.  Read Only.

Flow Meter Board 3		bd099	US bit	1 = Yes, 0 = No.  Read Only.

## JagBasic Application Interface to Remote Batching

ControlNet Class Code: 79 hex

- The Interface to Remote Batching Object has two instances, instances 1 and 2.

INTERFACE TO REMOTE BATCHING SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

(Instance 1 – PAC only contains variables 1-5 in instance 1)

1	Variable 1			bx101	US	Interface Variable to Remote Batching

Also, Auto-tune Setup State.

2	Variable 2			bx102	US	Interface Variable to Remote Batching

Also, Auto-tune Setup End Indication.

3	Variable 3			bx103	US	Interface Variable to Remote Batching

Also, Auto-tune Setup Error Counter.

4	Variable 4			bx104	US	Interface Variable to Remote Batching

Also, Auto-tune Setup Status.

5	Variable 5			bx105	US	Interface Variable to Remote Batching

Also, Auto-tune Setup Status Counter.

6	Variable 6			bx106	US	Interface Variable to Remote Batching

7	Variable 7			bx107	US	Interface Variable to Remote Batching

.	.				.	.	.

98	Variable 98			bx198	US	Interface Variable to Remote Batching

(Instance 2)

1	Variable 99			bx201	US	Interface Variable to Remote Batching

2	Variable 100			bx202	US	Interface Variable to Remote Batching

.	.				.	.	.

98	Variable 196			bx298	US	Interface Variable to Remote Batching

# SHARED DATA STATIC BRAM ELEMENTS

## Scale Weight Static Object

ControlNet Class Code: 66 hex

An instance of the Scale Weight Static Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SCALE WEIGHT STORED IN STATIC RAM SHARED DATA	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Scale Mode Out			wsn01	C	1 alphanumeric (GROSS or NET)

2	Displayed Tare Weight		wsn02	13C	12 alphanumeric,right-justified + 0

3	Displayed Aux Tare Weight	wsn03	13C	12 alphanumeric,right-justified + 0

4	Fine Tare Weight			wsn04	D	double float weight

5	Auxiliary Tare Weight		wsn05	D	double float weight

6	Current Units			wsn06	UC	1 = Primary,  2 = Secondary

7	Tare Source			wsn07	UC	1 = Pushbutton,  2 = Keyboard,  3 = Autotare

8	Current Zero Counts		wsn08	D	Double PB&AZM current zero counts

9	Tare Source String		wsn09	2C	“PT” = keyboard tare, else “T ”

10	Displayed Stored Weight		wsn10	13C	12 A/N, right justified.  Stores initial weight

for Net-Sign Correction.

11	Stored Weight			wsn11	D	Double float weight. Stores initial weight

for Net-Sign Correction.

12	Legal Tare Weight		wsn12	D	Double float weight.

13	Last Scale Error			wsn13	41C	Date - time - error message

14	Number Scale IO Errors		wsn14	F	Scale IO Errors since calibration or reset

- 20	Num Weighments Since Calibration	wsn20	F	Number of weighments since last calibration or reset.

- 21	Number Platform Overloads	wsn21	F	Number platform overloads since calibration or reset.

- 22	Number Platform High Impacts	wsn22	F	Number high impacts since calibration or reset.

- 23	Number Zero Commands		wsn23	F	Number zero commands since calibration or reset.

- 24	Number Zero Command Failures	wsn24	F	Number zero command failures due to “out of zero	range” since calibration or reset.

25	Symmetry Check Failure Type	wsn25	UC	0=No Failure

1= Estimate-able Symmetry Failure

2= Estimate-able Comm Failure

3= UnCorrectable Symmetry Failure

4= UnCorrectable Comm Failure

5= Estimate-able Zero Drift Failure

6= UnCorrectable Zero Drift Failure

- 26	RunFlat Detected Bad Cell		wsn26	UC	Power cell that was detected bad in symmetry check. 	If run flat is enabled, this cell is replaced using weight counts from replacement cell (0-23).

- 27	RunFlat Replacement Cell		wsn27	UC	Power cell that is used as replacement cell in run flat operation (0-23).

28	Calibration Check Failure		wsn28	UC	0 = No, 1 = Latest calibration check failed.

## Scale Calibration Object

ControlNet Class Code: 67 hex

An instance of the Scale Calibration Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SCALE CALIBRATION PARAMETERS STORED IN STATIC RAM 	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Auxiliary Display Units		csn01	UC	1=pounds, 2=kilograms, 3=grams,4=ounces,

5=pounds & ounces, 6=troy ounces, 7=penny weights, 8=metric tons, 9=tons, 10=custom units

2	Custom Units Name		csn02	7C	Custom Units Name.

3	Custom Units Conversion Factor	csn03	D	double float

4	Rate Time Units			csn04	C	1 alphanumeric (No, Sec, Min, Hour)

5	Rate Sample Period		csn05	UC	For csn06=0,1, and 2, this field specifies the number 	of seconds over which the	rate is averaged. For csn06=3, this field specifies the value for both	frequency and sample  period in number of weight updates.

6	Rate Calculation Frequency	csn06	UC	0 = every second, 1 = every five seconds,

2 = every half-second.

3 = calculation frequency and sample period

are specified in number of weight updates,

as specified in RateSampleTime.

7	WIM Mode			csn07	C	Analog Scale Weigh In Motion.  “W” = Yes.  Valid only for the first Single or Dual Analog Scale.

8	IDNET Higher Precision		csn08	C	Use IDNET Scale “times 10” Precision.

1 = Yes, 0 = No.

9	Power Up Timer			csn09	UC	2 alphanumeric, right-justified in minutes

10	Low-Pass Filter Corner Frequency	csn10	D	double float  (.1 Hz to 9.9 Hz by .1 Hz)

11	Notch Filter Frequency		csn11	D	double float  (.1 Hz to 9.9 Hz by .1 Hz)

12	Comb Filter Frequency		csn12	D	double float  (.1 Hz to 9.9 Hz by .1 Hz)

13	Print Threshold			csn13	D	double float weight

14	Print Reset Threshold		csn14	D	double float weight

15	Display Update Frequency		csn15	D	double float hertz

16	Custom Continuous Output Freq	csn16	D	double float hertz. Frequency for custom

continuous messages using Print Templates.

17	Low-Pass Filter Poles		csn17	US	unsigned integer.

18	Scale ID				csn18	20C	Scale Identifier Text.

19	Averaging Filter Order		csn19	US	unsigned integer.

20	Comb Filter Order		csn20	US	unsigned integer.

Scale Type			csn21	C	1 alphanumeric (Analog Load Cells, Power Digital Load Cells, IHigh Precision, Single cell DigiTOL, Power Module DigiTOL, UltraResHigh, or 	UltraResLow, SUmming )

ScaleLocation			csn22 	UC	0 = first unit, 1 = second unit (board or COM:

port)

IDNet Vibration Adapter		csn23	C	‘0’ - ‘9’ (specific to Precision Base)

IDNet Weighing Process Adapter	csn24	C	‘0’ - ‘9’ (specific to Precision Base)

IDNet Automatic Stability Detection	csn25	C	‘0’ - ‘9’ (specific to Precision Base)

IDNet Auto-Zero Setting		csn26	C	‘0’ = “Off”, ‘1’ = “On”

IDNet Software Part Number	csn27	12C	xxxx-x-xxxx string from Precision Base

IDNet Ident Code			csn28	2C	‘—’ to ‘99’ calibration count from Precision Base

Scale In Summing Scale		csn29	UC	Add Scale to Summing Scale,

0 = No, 1 = Yes

30	Calibration Date			csn30	12C	11 alphanumeric

85	Fillnoise Filter Enable		csn85	UC 	1 = True, 0 = False

86	Auto-Print			csn86	UC	1 = True, 0 = False

87	No Motion Before Print		csn87	UC	1 = True, 0 = False

88	Display Rate			csn88	UC	1 = True, 0 = False

89	Display Auxiliary Units		csn89	UC	1 = True, 0 = False

90	Units Switch Enable		csn90	UC	1 = True, 0 = False

91	Print Interlock Enable		csn91	UC	1 = True, 0 = False

92	Do IDNET Tare In Jag		csn92	UC	1 = True, 0 = False

93	Process Application		csn93	UC	0 = low, 1 = Mid, 2 = High

31	Next Scheduled Calibration Date 	csn31	L	In seconds since 1970 GMT

32	Calibration Interval In Days	csn32	US	Max number of days between calibrations

33	Calibration Interval In Weighments	csn33	L	Number of weighments between calibrations

34	Cal Expired Announcement	csn34	UC	1=log only, 2=disable scale and alarm, 3=email alert and alarm, 4=alarm only.

35	Last Calibration Date 		csn35	L	In seconds since 1970 GMT

36	Calibration Check Tolerance	csn36	D	Weight tolerance in primary units

37	Number Calibration Check-Points	csn37	UC	Number of calibration check points

38	Cal Check Failed Announcement	csn38	UC	1=log only, 2=disable scale and alarm, 3=email alertand alarm, 4=alarm only.

40	Monitor Cell Overloads		csn40	UC	0=No, 1=Count, 2=Count and Log

41	Cell Overload Threshold		csn41	D	Cell overload threshold in units in csn54

42	Monitor Platform Overload		csn42	UC	0,1=Count, 2=Count and Log

44	Monitor Platform High Impacts	csn44	UC	0=No, 1=Count, 2=Count and Log

45	High Impact Weight Threshold	csn45	D	Platform high impact weight in primary units

46	High Impact Rate Threshold	csn46	D	Platform high impact rate in primary units per second

47	Monitor Weighments		csn47	UC	0=No, 1=Count, 2=Count and Log

48	Weighment Trigger		csn48	UC	0=None, 1=Print Command, 2=Upscale Gross Weight Threshold, 3=Downscale Gross Weight Threshold,4=Upscale Net Weight Threshold, 5=Downscale

Net Weight Threshold

49	Weighment Threshold		csn49	D	Weighment threshold weight

50	Monitor Zero Commands		csn50	UC	0,1=Count, 2=Count and Log

51	Monitor Zero Command Failures	csn51	UC	0,1=Count, 2=Count and Log

52	Monitor Scale IO Errors		csn52	UC	0,1=Count, 2=Count and Log

53	Weighment Reset Threshold	csn53	D	Weighment reset threshold weight

54	Cell Overload Units		csn54	UC	1=counts, 2=primary units

55	Run Flat This Specific Cell		csn55	US	Use run flat on this specific cell

56	Threshold to Begin Symmetry	csn56	US	% of capacity to begin symmetry checking

RunTime Counter			cs157	D	Run Time counter in minutes.  Counts in run

mode when in normal weigh mode. (future)

Spare				cs257	D

Spare				cs357	D

Spare				cs457	D

Spare				cs557	D

Spare				csn58	D

Spare				csn59	D

NoWeightReadings		csn60	US	0=Valid wt. readings are occurring,

1=Valid wt. readings are not occurring

(future)

Spare				csn61	40C

62	Cell Symmetry Check		csn62	UC	0=No, 1=Count,2=Count and Log

63	Cell Zero Drift Check		csn63	UC	0=No, 1=Count,2=Count and Log

64	Cell Zero Drift Check Threshold	csn64	D	zero drift threshold in percent of span

65	Cell Symmetry Type		csn65	UC	0=No, 1=Radial, 2=Left-Right Pairs

66	Cell Symmetry Threshold		csn66	UC	percent difference (0-99)

67	Predictive Failure Announcement	csn67	UC	1=None, 2=disable scale and alarm, 3=email alert and alarm, 4=alarm only

68	RunFlat Weight Estimation		csn68	UC	0=No, 1=Yes

## Scale Tare Object

ControlNet Class Code: 68 hex

An instance of the Scale Tare Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SCALE TARE SHARED DATA	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Auto-Tare Threshold		trn01	D	double float  weight

2	Auto-Tare Reset Threshold	trn02	D	double float  weight

3	Auto-Clear Tare Threshold		trn03	D	double float  weight

85	Tare Enabled			trn85	UC	1 = True, 0 = False

86	Pushbutton Tare			trn86	UC	1 = True, 0 = False

87	Keyboard Tare			trn87	UC	1 = True, 0 = False

88	Auto-Tare			trn88	UC	1 = True, 0 = False

89	Auto-Tare Check Motion		trn89	UC	1 = True, 0 = False

90	Auto-Clear Tare			trn90	UC	1 = True, 0 = False

91	Auto-Clear Tare After Print		trn91	UC	1 = True, 0 = False

92	Auto-Clear Tare Motion		trn92	UC	1 = True, 0 = False

93	Tare Interlock			trn93	UC	1 = True, 0 = False

94	Display Tare			trn94	UC	1 = True, 0 = False

96	Net-Sign Correction		trn96	UC	1 = True, 0 = False

## Setpoint Object

ControlNet Class Code: 69 hex

An instance of the Setpoint Object may exist for each of 12 setpoints.  Each contains the numbered instance attributes.

SETPOINT SHARED DATA	(External Read/Write)

Where the ‘n’ is the instance representing the internal setpoint number (1-C). The default assignment for setpoint B is ZERO TOLERANCE for Scale A.  The default assignment for setpoint C is ZERO TOLERANCE for Scale B.

Attr. 	Function/Value			Name	Format	Description

1	Setpoint Name			spn01	9C	9 alphanumeric

2	Setpoint Enable Button		spn02	UC	1 alphanumeric (0=disabled, Scale 1,

Scale 2, 	Scale 3, Scale 4, Scale 5)

3	Setpoint Target Variable		spn03	C	1 alphanumeric

(Gross, Net, Displayed, Rate,

H=Gross With AutoPreact,

M=Net with AutoPreact,

L=LearnJog Weight Associated with a Timer Value,  J=Jog using Timer Value)

- 5	Setpoint Coincidence Value	spn05	D	Setup. For weight setpoint targets, this field

- has a weight value. For rate setpoints, this field contains the max value.  For LearnJog

- setpoints, this field contains a time value.

- 6	Setpoint Preact Value		spn06	D	Setup. For weight setpoint targets, this field

- is a cutoff preact value.  When this field is

- set, the setpoint turns off the feed when

- the  weight = spn05 – spn06. For AutoPreact setpoints, this field contains a time value in seconds.

- 8	Setpoint Dribble Value		spn08	D	Setup. For weight setpoint targets, this field

- is a feed dribble value.  When this field is

- set, the setpoint turns off the fast feed when

the  weight = spn05 – spn06 - spn08.

10	Setpoint Tolerance Value		spn10	D	Setup.  The Setpoint uses this field to

determine if the acutal cutoff weight falls

within this specified tolerance.

86	Setpoint Fill Or Discharge		spn86	UC	1 = Discharge, 0 = Fill

87	Setpoint Latching			spn87	UC	1 = Feed Latching Enabled.  Application

must set this field to enable “latching”.

When latching is set, the setpoint will

not re-enable the feed after the setpoint

coincidence is first reached until

PAC resets the “latched” bit.

88	Setpoint Latched			spn88	UC	1 = Latched, 0 = Unlatched.  If latching

(spn87 is set),  the setpoint sets this field

to 1 when it first encounters the setpoint

coincidence.  Application must reset this bit =

0 to start next setpoint processing.

When an application enables “Feed Latching”, the JagX O/S sets the Setpoint Latched=1 and the Setpoint Feeding=0 when it next encounters the setpoint coincidence.  Then, it never resets the Setpoint Feeding=0 condition again until the application resets the Setpoint Latched=0.  That is, the application must reset SetpointLatched=0 before starting a new setpoint.  Jog and LearnJog setpoints are always latched when the setpoints are started, so the application must reset the latch before starting a new setpoint.

The following four fields are secondary inputs to a single setpoint that has auto-adjusting preacts based

on flow rates. When the flow rate is greater than threshold 3, preact 3 is used. When the flow rate is greater

than threshold 2, preact 2 is used.  Threshold 3 is the higher threshold rate.  Threshold 2 is the lower

threshold rate.  If the flow rate is below both thresholds, the standard preact is used.

- Class Attributes

- Instance 0

Attr. 	Function/Value			Name	Format	Description

1	AutoAdjustSetpointThreshold2	spc08	D	double float weight

2	AutoAdjustSetpointPreact2		spc06	D	double float weight

3	AutoAdjustSetpointThreshold3	spb08	D	double float weight

4	AutoAdjustSetpointPreact3		spb06	D	double float weight

Jog Tables for the Jog setpoints are contained in the Cluster Variable fields.  The fields contain numbers

in string format.  Cluster variables 1-10 are the weight values.   Cluster Variables 11-20 are the timer

values associated with each of the weight values.  The weight and timer values must be ordered in

ascending order.  A weight value of 0 indicates the termination of the table values.

## System Object

ControlNet Class Code: 6A hex

- The System Object has only a single instance, instance 1.

SYSTEM SHARED DATA		(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	Current Selected Scale		jag01	3C	First Char =  “L” for local Node or

the Selected Remote Node =1-9, A-K.

2nd Char is Selected Scale = A, B, C, D, or E.

2	Selected Flow Meter Location	jag02	C	Run mode selected Flow Meter location = ‘L’

for local node or the selected remote node =

‘1’ – ‘9’, or ‘A’ – ‘K’.  See jag55 for Flow

Meter ID.

3	PAC Active Processes		jag03	6C	Currently active Q.iMPACT Processes.

Market				jag04	C	1 alphanumeric (USA, European Community, Australia, Canada)

5	Date Format			jag05	UC	1 byte integer

6	Time Format			jag06	UC	1 byte integer

7	Julian Date			jag07	9C	8 alphanumeric

8	Julian Time			jag08	9C	8 alphanumeric

9	Consecutive Number		jag09	L	long integer counter

10	Last Error Message		jag10	41C	Date - time - error message

Software ID			jag11	13C	12 alphanumeric

Spare				jag12	13C	12 alphanumeric

13	Number Discrete Inputs		jag13	UC	1 byte integer.  The number of discrete I/O’s

that are assigned as inputs, including 4 dedicated controller card discrete inputs plus the assigned MFIO inputs.  A value ‘B’ assigns default value 12 = 4 controller card inputs plus 8 MFIO inputs.

BRAM Version Number		jag14	L	4 byte integer.

Number Of Internal Scales		jag15	UC	1-5 Scales.

16	Date Separator			jag16	C	1 byte character.

17	Time Separator			jag17	C	1 byte character.

Consecutive Number Destination	jag18	11C	Size of J_FNAME + 1

19	Current Date	 		jag19	12C	11 alphanumeric

20	Time Of Day	 		jag20	12C	11 alphanumeric

21	Week Day	 		jag21	11C	10 alphanumeric

22	Consecutive Number Preset	jag22	L	CN Preset value

Character Set			jag23	UC	0=USA, 1=France, 2=England, 3=Germany,

4=Denmark-I, 5=Sweden, 6=Italy, 7=Spain-I,

8=Japan, 9=Norway, 10=Denmark-II, 11=Spain-II, 12=Latin America

Language			jag24	UC	0=English, 1=French, 2=German, 4=Spanish

Keyboard			jag25	UC	0=English, 1=French, 2=German, 4=Spanish

Scale Error Log Reset Time	jag26	24C	Date - time

Monitor Counts Reset Time	jag27	24C	Date - time

Monitor Log Reset Time		jag28	24C	Date - time

Monitor Log Max Records		jag29	L	Maximum number of records in Monitor Log File

Monitor Log Next Record Pointer	jag30	L	Pointer to next record in Monitor Log File

Monitor Log File Full Indicator 	jag31	UC	Indication that Monitor Log File is full and counter 	has been reset to beginning of file.

32	Last Demand Print Message 	jag32	80C	Last demand print message.

JagXtreme Serial #		jag33	20C	Mettler Toledo Serial # for this JagX.

JagXtreme ID			jag34	20C	User Assigned ID Name for this JagX.

JagXtreme Project		jag35	20C	User Assigned Project Name for this JagX.

JagXtreme Description		jag36	80C	User Assigned Description for this JagX.

Indirect Read Counter		jag37	US	Counter for Reading Shared Data Sequentially.

Power Cycle Counter		jag38	F	Power Cycle Counter.

SD Indirect Read Pointer 11	jag39	5C	See description for jag40

SD Indirect Read Pointer 1		jag40	5C	This field points to another Shared Data field. The Web Pages may set this field to point to another Shared Data field in order to read the other field through the Web Server Side Includes.

SD Indirect Read Pointer 2		jag41	5C	See description for jag40

SD Indirect Read Pointer 3		jag42	5C	See description for jag40

SD Indirect Read Pointer 4		jag43	5C	See description for jag40

SD Indirect Read Pointer 5		jag44	5C	See description for jag40

SD Indirect Read Pointer 6		jag45	5C	See description for jag40

SD Indirect Read Pointer 7		jag46	5C	See description for jag40

SD Indirect Read Pointer 8		jag47	5C	See description for jag40

SD Indirect Read Pointer 9		jag48	5C	See description for jag40

SD Indirect Read Pointer 10	jag49	5C	See description for jag40

- Class 0x04 Assembly Object, Instance 0

- Attr.

100	Configurable Assembly 1		jag50	50C	This Shared Data field contains a list of up to 10 Shared Data fields.  This list enables a remote PLC to read multiple Shared Data fields concurrently.

101	Configurable Assembly 2		jag51	50C	See description of jag50.

102	Configurable Assembly 3		jag52	50C	See description of jag50.

103	Configurable Assembly 4		jag53	50C	See description of jag50.

Flow Meter Selected in Setup	jag54	UC	Currently selected Flow Meter in Setup

Mode.

Flow Meter Selected in Run Mode	jag55	C	Run mode selected Flow Meter = ‘K’ – ‘V’.

See jag02 for Flow Meter location.

Mode Change Request		jag56	UC	Request Mode Change from Web browser

1=Enter Setup req., 2=Exit Setup req.,

cleared when mode request is processed.

Web Page Language		jag57	US 	0=English, 1= Non-English

Number of Flow Meters		jag58	US	Number of internal Flow Meters (0-12).

Console Print Server Routing	jag59	US	Bit field for routing data to Con. Print Server:

Bit 0 = 1, Route JagBASIC LPRINTs to the

Console Print Server port 1749.

Bit 1 = 1, Route maintenance reports to the

Console Print Server port 1749.

Bit 2 = 1, Route terminal setup report to the

Console Print Server port 1749. (future)

HMI Subscription String		jag60	10C	The HMI can request specific streams of scale, lower JagX display, and JagBasic message displays.  The string contains a subset of <ABCDESJLT>, where ABCDE represents the scales, S represents the

selected scale, L represents the lower JagX

Display,	 J is the JagBasic Message Field,

and T is time.  See description of bd014.

Alibi Memory-Max Records	jag61	L	Maximum number of records in Alibi Memory

File

Alibi Memory-Next Record Pointer	jag62	L	Pointer to next record in Alibi Memory File

Alibi Memory-File Full Indicator 	jag63	UC	Indication that Alibi Memory File is full and

counter 	has been reset to beginning of file.

Alibi Memory-Post Request	jag64	40C	Search string for locating a particular

transaction record in Alibi Memory File.

Battery Replacement Text		jag65	40C	Battery replacement text string.

PAC Scale ID, Flow Meter ID	jag66	3C	The first two bytes are The terminal number

(01-20) and the last Byte is the scale letter

(A-D) or flow meter (K-V).  When the Web

Browser reads the alias bd054, the Web

Server returns a string:

“uuu^aaaaaaaa^bbbbbbbb^P^iiiiiiiiiiiiiiiiiiii”

where uuu is  the units {“ lb”,” kg“,”  g“, or “

t“}, aaaaaaa is the increment size, bbbbbbb

is the scale capacity, P is the rate period

{“N”, “H”,	“M”, or “S” }, and iiiiiiiiiiiiiiiiiiii is the

measuring device identifier.

Terminal Download Status		jag67	4C	Download of shared data to terminal status

1st two chars = terminal no. 3rd char status

‘S’ = Success, ‘F’=Fail

85	Key Beeper Enable		jag85	UC 	1 = On,  0 = Off

86	Alarm Beeper Enable		jag86	UC	1 = On,  0 = Off

Auto Configuration		jag87	UC	0=AutoConfigure Scales, 4=Receive Literals,

5=Receive BRAM, 6=Receive EEPROM,

14=Send Literals, 15=Send BRAM,

16= Send EEPROM

88	Legal For Trade			jag88	UC	1 = True,  0 = False

89	Consecutive Number Enable	jag89	UC	1 = True,  0 = False

90	Consecutive Number Preset Enable	jag90	UC	1 = True,  0 = False

91	Disable Memory Key		jag91	UC	1 = True,  0 = False

Alibi Memory-Enabled		jag92	UC	1 = True,  0 = False

## User Literals Object

ControlNet Class Code: 6B hex

- The User Literals Object has only a single instance, instance 1.

USER LITERALS SHARED DATA	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

1	User Literal 1			lit01	40C	40 alphanumeric

2	User Literal 2			lit02	40C	40 alphanumeric

*			  	*	    *		*

21	User Literal 21	 		lit21	40C	40 alphanumeric

*   				  *	  *                         *

50	User Literal 50	 		lit50	40C	40 alphanumeric

## User Prompts Object

ControlNet Class Code: 6C hex

- The User Prompts Object has only a single instance, instance 1.

USER PROMPTS SHARED DATA	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

1	User Prompt 1			pmt01	16C	16 alphanumeric

2	User Prompt 2			pmt02	16C	16 alphanumeric

*			  	*	 *	*

20	User Prompt 20			pmt20	16C	16 alphanumeric

## User Variables Object

ControlNet Class Code: 6D hex

- The User Variables Object has only a single instance, instance 1.

USER VARIABLES SHARED DATA	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

1	User Variable 1			var01	47C	USER_VARIABLE structure

2	User Variable 2			var02	47C	USER_VARIABLE structure

*			  	*	   *		*

20	User Variable 20			var20	47C	USER_VARIABLE structure

81	Variables In Use			var81	UC	Number 0-20

82	Prompt Looping Mode		var82	UC	0 = No Loop,   1 = Loop

## Cluster Variables Object

ControlNet Class Code: 6E hex

- The Cluster Variables Object has only a single instance, instance 1.

CLUSTER VARIABLES SHARED DATA	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

1	Cluster Variable 1			clv01	40C	40 alphanumeric

2	Cluster Variable 2			clv02	40C	40 alphanumeric

*			 	 *	     *					*

20	Cluster Variable 20		clv20	40C	40 alphanumeric

Cluster Variable fields may contain Jog Tables for the Jog setpoints.  The fields have numbers in string

format.  Cluster variables 1-10 are the weight values.  Cluster Variables 11-20 are the associated timer

values.

## Print Templates

TEMPLATES SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Printer Template 1		ptp01	489C	480 a/n grammer+8 a/n template name+0

Printer Template 2		ptp02	489C	480 a/n grammer+8 a/n template name+0

Printer Template 3		ptp03	489C	480 a/n grammer+8 a/n template name+0

Printer Template 4		ptp04	489C	480 a/n grammer+8 a/n template name+0

Printer Template 5		ptp05	489C	480 a/n grammer+8 a/n template name+0

## Security

SECURITY SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

SecurityScaleInterface		scu01	C	Reserved for JagX O/S Use Only.

1 alphanumeric (  None, Super, Maint,LegalforTrade)

SecurityApplicationEnviron		scu02	C	Reserved for JagX O/S Use Only.

SecuritySerialInterface		scu03	C	Reserved for JagX O/S Use Only.

SecurityDiscreteConfig		scu04	C	Reserved for JagX O/S Use Only.

SecurityMemoryConfig		scu05	C	Reserved for JagX O/S Use Only.

SecurityJagBasicConfig		scu06	C	Reserved for JagX O/S Use Only.

SecurityNetworkConfig		scu07	C	Reserved for JagX O/S Use Only.

SecuityDiagnostics		scu08	C	Reserved for JagX O/S Use Only.

SecurityAllenBradley		scu09	C	Reserved for JagX O/S Use Only.

SecurityUnusedForNow		scu10	C	Reserved for JagX O/S Use Only.

SupervisorPassword		scu11	9C	Reserved for JagX O/S Use Only.

MaintenancePassword		scu12	9C	Reserved for JagX O/S Use Only.

## Serial Ports

SERIAL PORT SETUP SHARED DATA	(External Read Only)

Where the ‘n’ is the serial port number (1-4).

Attr. 	Function/Value			Name	Format	Description

Interface Type:			srn01	UC	0=RS232, 1=RS422, 2=RS485

Baud Rate			srn02	UC	0=300, 1=600, 2=1200, 3=2400, 4=4800, 5=9600,6=19200,7=38400,8=57600,9=76800,10=115200.

Reserved			srn03	UC	Reserved for JagX O/S Use Only.

Parity				srn04	UC	Same as BIOS values. 0=even, 16=odd,

64=none

Flow Control			srn05	UC	Same as BIOS values. 0=none, 1=Xon/Xoff,

2=RS232.

Input Connection			srn06	FPTR	Reserved for JagX O/S Use Only.

Data Bits			srn07	UC	Same as BIOS values. 8=7 bits, 12=8 bits

Stop Bits			srn08	UC	Same as BIOS values. 1=1, 2=1.5, 3=2

Owner Name			srn09	7C	Reserved for JagX O/S Use Only.

Checksum			srn85	UC	1 = On, 0 = Off

## Network

NETWORK INTERFACE SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Ethernet Node Address		net01	13C	Ethernet Node Address

Default Ethernet IP Address	net02	UL	Internet Protocol Address

Ethernet IP Address Subnet Mask	net03	UL	Node ID Address Subnet Mask

Ethernet Gateway IP Address	net04	UL	Default Gateway IP Address

Ethernet DNS Server IP Address 1	net05	UL	Domain Name Server (DNS) IP Address 1

Ethernet DNS Server IP Address 2	net06	UL	Domain Name Server (DNS) IP Address 2

Ethernet DNS Server IP Address 3	net07	UL	Domain Name Server (DNS) IP Address 3

MAX Cluster Nodes		net08	UC	Maximum Number of Nodes

Remote Cluster Node Count	net09	UC	Remote Node Count

Cluster Node Number of this Node 	net10	UC	Cluster Node Number of this node –

1,2,3,4,5,or 6.

PPP User Name			net11	20C	Local (Point-to-Point Protocol) PPP Login

Name

PPP User Password		net12	20C	Local PPP User Password

PPP Dial Phone Number		net13	32C	PPP Phone Number

PPP Modem AT Initialization 	net14	32C	PPP Modem Initialization String

PPP IP Address			net15	UL	PPP Internet Protocol Address

PPP IP Address Subnet Mask	net16	UL	PPP Node ID Address Subnet Mask

PPP Gateway IP Address		net17	UL	PPP Default Gateway IP Address

PPP Keep Alive			net18	US	Seconds between PPP echoes. 0 = disable

PPP Client Timeout		net19	US	Timeout for Connects as Client

PPP CHAP Protocol Secret	net20	20C	Secret for Chap/MD5 Hashing

PPP Login File			net21	20C	Remote PPP Login Name

PPP Log Server File		net22	20C	Remote PPP User Password

SMTP Server IP Address		net23	UL	SMTP (E-mail) Server IP Address

SMTP Sending Machine Name	net24	40C	SMTP (E-mail) Sending Machine Name

SMTP Sender E-mail Address	net25	40C	SMTP (E-mail) Sender’s Email Address

SMTP Subject			net26	40C	SMTP (E-mail) Subject Line

SMTP IO Port			net27	UC	1 = Ethernet, 2 = PPP

SMTP Server TCP Port		net28	US	SMTP Server TCP Port.  Default = 25.

Home Page			net30	80C	Home Page for Web Server

Documentation Page		net31	80C	Documentation Page for Web Server

Help Page			net32	80C	Help Page for Web Server

Ethernet DHCP Client		net85	UC	1=Yes.  Use Dynamic Host Configuration

Protocol 	Client to get this node’s Ethernet IP

Address. 0=No. Use IP Address in Shared

Data field  “net02”.

Enable DNS Client		net86	UC	1=Yes.   Use Domain Name System Servers

to find IP Addresses of cluster nodes and

hosts. 0=No. Use IP Addresses in Shared

Data fields rmn00 and rwn00.

Enable Web Server		net87	UC	1=Yes, 0=No

Enable FTP Server		net88	UC	1=Yes, 0=No; File Transfer Protocol.

Enable Email Alert		net89	UC	1=Yes, 0=No

Network Console			net91	UC	1 = Yes, 0 = No

Enable PPP Auto-Dial		net92	UC	1=Yes, 0=No

Enable PPP Auto-Answer		net93	UC	1=Yes, 0=No

PPP Console Log			net94	UC	1=Yes, 0=No

PPP Require CHAP		net96	UC	1=Yes, 0=No; Challenge Handshake

Authentication Protocol.

PPP Require PAP		net97	UC	1=Yes, 0=No; Password Authentication

Protocol.

PPP Log to LPRINT Device	net98	UC	1=Yes, 0=No;

## Email Alert

EMAIL ALERT RECIPIENT SHARED DATA	(External Read Only)

Where the ‘n’ is an index number (1-6).

Attr. 	Function/Value			Name	Format	Description

E-mail Address			emn01	40C	E-mail Address of Recipient

E-mail Calibration Checks		emn02 	UC	1=Yes, 0=No

E-mail Predictive Failures		emn03 	UC	1=Yes, 0=No

Spare				emn04 	UC

Spare				emn05 	UC

## FTP Login

Entries 1-6 are used for FTP user name/passwords.  Entries 7-9 are spares Reserved for future use.

FTP LOGIN SHARED DATA	(External Read Only)

Valid user names, passwords for logging into the JagX where the ‘n’ is an index number (1-9).

Attr. 	Function/Value			Name	Format	Description

Login User Name			lgn01	20C	Login User Name

Login Password 			lgn02	20C	Login Password

Ethernet/IP IP Address		lg801	20C	Setup.  IP Address for ControlNet over

Ethernet in Q.iMPACT Rev. B only.

Ethernet/IP Subnet Mask		lg802	20C	Setup. Subnet mask for ControlNet over

Ethernet in Q.iMPACT Rev. B only.

Ethernet/IP Gateway Address	lg901	20C	Setup.  Gateway address for ControlNet

over Ethernet in Q.iMPACT Rev. B only.

Ethernet/IP DHCP Client		lg902	20C	Setup. DHCP Client for ControlNet over

Ethernet in Q.iMPACT in Rev. B only.

‘1’=Y, ‘0’=N

## Network Remote Nodes

### Remote Node Addresses

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Cluster Node Name _ 1		nde01	30C	Not used in Q.iMPACT.

Cluster Node Name _ 2		nde02	30C	Not used in Q.iMPACT.

Cluster Node Name _ 3		nde03	30C	Not used in Q.iMPACT.

Cluster Node Name _ 4		nde04	30C	Not used in Q.iMPACT.

Cluster Node Name _ 5		nde05	30C	Not used in Q.iMPACT.

Cluster Node Name _ 6		nde06	30C	Not used in Q.iMPACT.

Host Node Name _ 1		nde07	30C	Not used in Q.iMPACT.

Host Node Name _ 2		nde08	30C	Not used in Q.iMPACT.

PPP Host Node Name 		nde09	30C	Not used in Q.iMPACT.

Cluster Node IP Address _ 1	nde11	UL	Refer to Q.iMPACT definition.

Cluster Node IP Address _ 2	nde12	UL	Refer to Q.iMPACT definition.

Cluster Node IP Address _ 3	nde13	UL	Refer to Q.iMPACT definition.

Cluster Node IP Address _ 4	nde14	UL	Refer to Q.iMPACT definition.

Cluster Node IP Address _ 5	nde15	UL	Refer to Q.iMPACT definition.

Cluster Node IP Address _ 6	nde16	UL	Refer to Q.iMPACT definition.

Host Node IP Address _ 1		nde17	UL	Refer to Q.iMPACT definition.

Host Node IP Address _ 2		nde18	UL	Refer to Q.iMPACT definition.

PPP Host Node IP Address	nde19	UL	Refer to Q.iMPACT definition.

### Remote Node Enable

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

ClusterConnection_1		rm101	UC	Refer to Q.iMPACT definition.

ClusterConnection_2		rm102	UC	Refer to Q.iMPACT definition.

ClusterConnection_3		rm103	UC	Refer to Q.iMPACT definition.

ClusterConnection_4		rm104	UC	Refer to Q.iMPACT definition.

ClusterConnection_5		rm105	UC	Refer to Q.iMPACT definition.

ClusterConnection_6		rm106	UC	Refer to Q.iMPACT definition.

HostConnection_1		rm107	UC	Refer to Q.iMPACT definition.

HostConnection_2		rm108	UC	Refer to Q.iMPACT definition.

PPP Connection			rm109	UC	Refer to Q.iMPACT definition.

## PLC Configuration

### PLC Configuration Shared Data

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Rack Address			abc01	UC	Allen-Bradley 0-59,

PROFIBUS station ID 1-127,

Modbus Plus 1-63.

Allen-Bradley Starting Quarter	abc02	UC	1 - 4

Allen-Bradley Data Rate		abc03	UC	0 = 57.6k, 1 = 115.2k,2 = 230.4k

Number Of Setpoints		abc04	UC 	Reserved for JagX O/S Use Only.

Number Of Scales		abc05	UC	1-4

Discrete Data Format		abc06	UC	0=Weight, 1=Increments, 2=Extended

Weight, 3=Floating Point

Input Rotation			abc07	10C	10 character string

Allen-Bradley Last Rack		abc85	UC	1 = Yes,  0 = No

85	ControlNet Media Layer		abc85	UC	Setup. 1= ControlNet, 2 = EtherNet/IP

Block Transfer Enable		abc86	UC	1 = Yes,  0 = No

Modbus Plus Globals Enable	abc87	UC	1 = Yes,  0 = No

PLC Controls Scale A Setpoints	abc88	UC	1 = Yes,  0 = No

PLC Controls Scale B Setpoints	abc89	UC	1 = Yes,  0 = No

### PLC (Non-PAC) Scale Configuration

(External Read Only)

Where the ‘n’ is a PLC scale index number (1-4).

Attr. 	Function/Value			Name	Format	Description

Terminal Node Name		abn01	2C	2 alphanumeric (J1,J2,J3,J4,J5,J6)

Scale Selection			abn02	UC	1 byte unsigned integer (1,2,3,4,5)

Scale Location			abn85	UC	0 = Local, 1 = Remote

### Analog Output

(External Read Only)

Where the ‘n’ is a channel index number (1-2).

Attr. 	Function/Value			Name	Format	Description

Analog Out Source Data		aon01	C	G=Gross Weight Scale 1,	H=Gross Weight

Scale 2,	I=Gross Weight Scale 3,	J=Gross

Weight Scale 4,	K=Gross Weight Scale 5

L=Net Weight Scale 1,

M=Net Weight Scale 2,

N=Net Weight Scale 3,

O=Net Weight Scale 4,

P=Net Weight Scale 5

Q=Rate Scale 1, R=Rate Scale 2,

S=Rate Scale 3,

T=Scale 4,U=Rate Scale 5

B=JagBasic Scale 1, C=JagBasic Scale 2.

D=JagBasic Scale 3, E=JagBasic Scale 4,

F=JagBasic Scale 5.

Analog Out Adjusted Zero		aon02	D	Zero Adjustment Offset

Analog Out AdjustedSpan		aon03	D	Full Scale Adjustment Offset

Analog Out Zero Preset		aon04	D	Zero Adjustment Offset

Analog Out Span Preset		aon05	D	Full Scale Adjustment Offset

## CONNECTION SHARED DATA

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Connection Object 1		cnn01	26C	Structure. Reserved for JagX O/S Use.

Connection Object 2		cnn02	26C	Structure. Reserved for JagX O/S Use.

*				    *

Connection Object 12		cnn12	26C	Structure. Reserved for JagX O/S Use.

## LADDER LOGIC DATA

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

Lader Rung Counter		lad01	US	Number of active rungs in ladder.

Ladder Rungs			lad02	400C	Ladder Structure, containing up to 50 rungs.

Each rung is 8 bytes.

## EEPROM SAVE AREA

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

EEPROM Save Area A		eb101	290UC	Reserved for JagX O/S Use Only.

EEPROM Write In Progress A	eb185	US bit	Reserved for JagX O/S Use Only.

EEPROM Save Area B		eb201	290UC	Reserved for JagX O/S Use Only.

EEPROM Write In Progress B	eb285	US bit	Reserved for JagX O/S Use Only.

EEPROM Save Area C		eb301	290UC	Reserved for JagX O/S Use Only.

EEPROM Write In Progress C	eb385	US bit	Reserved for JagX O/S Use Only.

EEPROM Save Area D		eb401	290UC	Reserved for JagX O/S Use Only.

EEPROM Write In Progress D	eb485	US bit	Reserved for JagX O/S Use Only.

## BASIC Application Object

ControlNet Class Code: 70 hex

- The BASIC Application Object has only a single instance, instance 1.

BASIC Application Shared Data	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

Program 1			bas01	20C	19 Alphanumeric characters + 0

Program 2			bas02	20C	19 Alphanumeric characters + 0

Program 3			bas03	20C	19 Alphanumeric characters + 0

Program 4			bas04	20C	19 Alphanumeric characters + 0

Program 5			bas05	20C	19 Alphanumeric characters + 0

Program 6			bas06	20C	19 Alphanumeric characters + 0

Program 7			bas07	20C	19 Alphanumeric characters + 0

Program 8			bas08	20C	19 Alphanumeric characters + 0

Program 9			bas09	20C	19 Alphanumeric characters + 0

Keyboard Source			bas10	UC	0=None, 1=Keypad, 2=Keyboard, 3=Both

Display Destination		bas11	UC	0=None, 1=Lower Display, 2=Serial Port

12	Programmable Tare Scale A	bas12	D	double float weight

13	Programmable Tare Scale B	bas13	D	double float weight

14	CustomOutput_A1_FromPLC	bas14	F	Float.	Defined by user.  Scale A.

15	CustomOutput_A2_FromPLC	bas15	4C	String.	Defined by user.  Scale A.

16	CustomOutput_A3_FromPLC	bas16	F	Float.   	Defined by user.  Scale A.

17	CustomOutput_A4_FromPLC	bas17	4C	String.  	Defined by user.  Scale A.

18	CustomInput_A1_ToPLC		bas18	F	Float.   	Defined by user.  Scale A.  High-

Speed.

19	CustomInput_A2_ToPLC		bas19	4C	String.  	Defined by user.  Scale A.  High-

Speed.

20	CustomInput_A3_ToPLC		bas20	F	Float.   	Defined by user.  Scale A.

21	CustomInput_A4_ToPLC		bas21	4C	String.  	Defined by user.  Scale A.

22	CustomOutput_B1_FromPLC	bas22	F	Float. 	Defined by user.  Scale B.

23	CustomOutput_B2_FromPLC	bas23	4C	String. 	Defined by user.  Scale B.

24	CustomOutput_B3_FromPLC	bas24	F	Float. 	Defined by user.  Scale B.

25	CustomOutput_B4_FromPLC	bas25	4C	String. 	Defined by user.  Scale B.

- 26	CustomInput_B1_ToPLC		bas26	F	Float. 	Defined by user.  Scale B.  High-Speed.

- 27	CustomInput_B2_ToPLC		bas27	4C	String. 	Defined by user.  Scale B.  High-Speed.

28	CustomInput_B3_ToPLC		bas28	F	Float. 	Defined by user.  Scale B.

29	CustomInput_B4_ToPLC		bas29	4C	String. 	Defined by user.  Scale B.

30	CustomOutput_C1_FromPLC	bas30	F	Float.	Defined by user.  Scale C.

31	CustomOutput_C2_FromPLC	bas31	4C	String.	Defined by user.  Scale C.

32	CustomOutput_C3_FromPLC	bas32	F	Float.   	Defined by user.  Scale C.

33	CustomOutput_C4_FromPLC	bas33	4C	String.  	Defined by user.  Scale C.

34	CustomInput_C1_ToPLC		bas34	F	Float.   	Defined by user.  Scale C.  High-

Speed.

35	CustomInput_C2_ToPLC		bas35	4C	String.  	Defined by user.  Scale C.  High-

Speed.

36	CustomInput_C3_ToPLC		bas36	F	Float.   	Defined by user.  Scale C.

37	CustomInput_C4_ToPLC		bas37	4C	String.  	Defined by user.  Scale C.

38	CustomOutput_D1_FromPLC	bas38	F	Float. 	Defined by user.  Scale D.

39	CustomOutput_D2_FromPLC	bas39	4C	String. 	Defined by user.  Scale D.

40	CustomOutput_D3_FromPLC	bas40	F	Float. 	Defined by user.  Scale D.

41	CustomOutput_D4_FromPLC	bas41	4C	String. 	Defined by user.  Scale D.

42	CustomInput_D1_ToPLC		bas42	F	Float. 	Defined by user.  Scale D.  High-

Speed.

43	CustomInput_D2_ToPLC		bas43	4C	String. 	Defined by user.  Scale D.  High-

Speed.

44	CustomInput_D3_ToPLC		bas44	F	Float. 	Defined by user.  Scale D.

54	CustomInput_D4_ToPLC		bas45	4C	String. 	Defined by user.  Scale D.

46	CustomOutput_E1_FromPLC	bas46	F	Float. 	Defined by user.  Scale E.

47	CustomOutput_E2_FromPLC	bas47	4C	String. 	Defined by user.  Scale E.

48	CustomOutput_E3_FromPLC	bas48	F	Float. 	Defined by user.  Scale E.

49	CustomOutput_E4_FromPLC	bas49	4C	String. 	Defined by user.  Scale E.

50	CustomInput_E1_ToPLC		bas50	F	Float. 	Defined by user.  Scale E.  High-

Speed.

51	CustomInput_E2_ToPLC		bas51	4C	String. 	Defined by user.  Scale E.  High-

Speed.

52	CustomInput_E3_ToPLC		bas52	F	Float. 	Defined by user.  Scale E.

53	CustomInput_E4_ToPLC		bas53	4C	String. 	Defined by user.  Scale E.

54	Programmable Tare Scale C	bas54	D	double float weight

55	Programmable Tare Scale D	bas55	D	double float weight

56	Programmable Tare Scale E	bas56	D	double float weight

Auto Start Enabled		bas85	UC	1 = True, 0 = False

Escape Enabled			bas86	UC	1 = True, 0 = False

Select Enabled			bas87	UC	1 = True, 0 = False

Manual Start Enabled		bas88	UC	1 = True, 0 = False

Manual Stop Enabled		bas89	UC	1 = True, 0 = False

## Power Cell Log Object

ControlNet Class Code: 71 hex

- The Power Cell Object has only a single instance, instance 1.

POWER CELL LOG		(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	Number IO Errors-Cell 1-24	pcl01	24L 	96 bytes. One long integer for each cell.

3	Current Zero Counts-Cell 1-24	pcl03	24L 	96 bytes. One long integer for each cell.

5	Number Cell Overloads-Cell 1-24	pcl05	24L	96 bytes. One long integer for each cell.

6	Num Symmetry Failures-Cell 1-24	pcl06	24L	96 bytes. One long integer for each cell.

7	Num Zero Drift Failures-Cell 1-24	pcl07	24L	96 bytes. One long integer for each cell.

# SHARED DATA EEPROM ELEMENTS

## Scale Calibration - EEPROM Object

ControlNet Class Code: 72 hex

An instance of the Scale Calibration – EEPROM Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SCALE CALIBRATION PARAMETERS STORED IN EEPROM	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

Address Of First Load Cell		cen01	UC	Power Cell starting address

Number Load Cells		cen02	UC	unsigned 0-255

3	Primary Units			cen03	UC	1 alphanumeric (1=pounds,2=kilograms,

3=grams, 4=metric tons)

4	Primary Number Ranges		cen04	UC	1 alphanumeric

5	Primary Low Increment Size	cen05	D	double float weight

6	Primary Mid Increment Size	cen06	D	double float weight

7	Primary High Increment Size	cen07	D	double float weight

8	Primary LowMid Threshold		cen08	D	double float weight

9	Primary MidHigh Threshold	cen09	D	double float weight

10	Primary Scale Capacity		cen10	D	double float weight

11	Secondary Units			cen11	UC	1 alphanumeric (1=pounds, 2=kilograms,

3=grams, 4=metric tons)

12	Secondary Number Ranges	cen12	UC	1 alphanumeric

13	Secondary Low Increment Size	cen13	D	double float weight

14	Secondary Mid Increment Size	cen14	D	double float weight

15	Secondary High Increment Size	cen15	D	double float weight

16	Secondary LowMid Threshold	cen16	D	double float weight

17	Secondary MidHigh Threshold	cen17	D	double float weight

18	Secondary Scale Capacity		cen18	D	double float weight

19	Calibration Units			cen19	UC	1 alphanumeric (1=primary or 2=secondary)

20	Zero Calibration Counts		cen20	L	integer

21	High Calibration Counts		cen21	L	integer

22	High Calibration Weight		cen22	D	double float weight

23	Mid Calibration Counts		cen23	L	integer

24	Mid Calibration Weight		cen24	D	double float weight

25	Gravity Adjust			cen25	D	double float

26	Motion Stability Sensitivity in D	cen26	F	float divisions

27	Motion Stability Time Period	cen27	UC	Time in tenths of seconds

28	ScaleSerialNumber		cen28	13C	12 alphanumeric

29	Calibration Counter 1		cen29	UC	Calibration Check Counter

30	Calibration Counter 2		cen30	UC	Calibration Counter

31	AtoD Update Rate		cen31	UC	Conversions per Second (1-255)

32	Over Capacity Divisions		cen32	UC	number of divisions (1-255)

85	Linearity Correction Enable	cen85	US bit	1 = True,  0 = False

86	Over Capacity Blanking		cen86	US bit	1 = True,  0 = False

87	Multirange Mode			cen87	US bit	1 = Auto,  0 = Manual

88	Shift Adjust Mode			cen88	US bit	0 = Cell,  1 = Pair

## EEPROM Version Identification

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

EEPROM Version Id (Scale A)	ee101	L	Long Integer (32 bits)

EEPROM Version Id (Scale B)	ee201	L	Long Integer (32 bits)

## Shift Adjust Object

ControlNet Class Code: 73 hex

An instance of the Shift Adjust Object may exist for each of 5 scales.  Each contains the numbered instance attributes.

SHIFT ADJUST VARIABLES	(External Read Only)

Where the ‘n’ is the instance representing the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Cell #1 Shift Constants		san01	L	Long Integer (32 bits) Normalized

“			   “		        “

16	Cell #16 Shift Constants		san16	L	Long Integer (32 bits) Normalized

17	Cell #17 Shift Constants		sxn17	L	Long Integer (32 bits) Normalized

“			   “	 “ 	        “

24	Cell #24 Shift Constants		sxn24	L	Long Integer (32 bits) Normalized

## Cell Calibration Object

ControlNet Class Code: 74 hex

- The Cell Calibration Object has only a single instance, instance 1.

INDIVIDUAL CELL CALIBRATION COUNTS  (External Read Only)

Where the ‘n’ is the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

1	Calibrated Zero Counts-Cell 1-24	ccn01	24L 	96 bytes. One long integer for each cell.

2	Calibrated  Span Counts-Cell 1-24	ccn02	24L	96 bytes. One long integer for each cell.

## Scale Zero

(External Read Only)

Where the ‘n’ is the internal scale number 1-5.

Attr. 	Function/Value			Name	Format	Description

Power-Up Zero Capture Pos Range	zrn01	UC 	percent capacity  (0 - 99)

Power-Up Zero Capture Neg Range	zrn02	UC	percent capacity  (0 - 99)

Pushbutton Zero Positive Range	zrn03	UC	percent capacity  (0 - 99)

Pushbutton Zero Negative Range	zrn04	UC 	percent capacity  (0 - 99)

Auto-Zero Maintenance Window 	zrn05	F	float - number of divisions

Under-Zero Divisions		zrn06	UC	0-99 divisions.  Number of divisions at which

the under-zero indication is turned on the

the display.  “99” disables then under-zero

display.

Pushbutton Zero			zrn85	US bit	1 = True, 0 = False

Auto-Zero in Gross Mode		zrn86	US bit	1 = True, 0 = False

Auto-Zero in Gross&Net Mode	zrn87	US bit	1 = True, 0 = False

Zero-Indication in Gross Mode	zrn88	US bit	1 = True, 0 = False

Zero-Indication in Gross&Net Mode	zrn89	US bit	1 = True, 0 = False

# GLOBAL DISCRETE I/O DATA

Global Discrete I/O data are discrete bit fields representing both physical I/O and logical I/O, The logical I/O may have either level sensitive states or edge sensitive events. Level sensitive fields can generate events when either a 0 or a 1 is written to the field. Edge sensitive bit fields only trigger events when a 1 is written to the field.  Global Discrete I/O is transitory and is not saved across a power down. On power up it is initialized to an all zero state.

- Note: All ControlNet Attribute Numbers from this point forward are in hexadecimal format.

## Level Sensitive Discrete Status Object

ControlNet Class Code: 75 hex

- The Level Sensitive Discrete Status Object has only a single instance, instance 1.

- NOTE: This object’s attributes are offset by 200hex (0x200); Actual Data = (0x200 + (Attr. No. –1))

some t_600 values are also contained at the end of this object and are offset 0x600.

- LEVEL SENSITIVE LOGICAL DISCRETE I/O DATA

Attr. 	Function/Value			Name	Format	Description

### Status of Scale A

01	Motion Out A			s_200	US bit	1 = True, 0 = False.  Read only.

02	Center Of Zero A			s_201	US bit	1 = True, 0 = False.  Read only.

03	Over Capacity A			s_202	US bit	1 = True, 0 = False.  Read only.

04	Under Zero A			s_203	US bit	1 = True, 0 = False.  Read only.

05	Net Mode A			s_204	US bit	1 = True, 0 = False.  Read only.

06	Scale Critical Error A		s_205	US bit	1 = True, 0 = False.  Read only.

07	Stored Weight Mode A		s_206	US bit	1 = True, 0 = False.  Read only.

08	Scale Selected A			s_207	US bit	1 = True, 0 = False.  Read only.

IDNET In Motion Error A		s_260 	US bit	Reserved for JagX O/S Use Only.

62	Weight Data OK A		s_261	US bit	1 = True, 0 = False.  Read Only.

63	Rate Setpoint OK A		s_262	US bit	1 = True, 0 = False.  Read Only.

64	Estimated Weight _A		s_263	US bit	1 = True, 0 = False.  Read Only.

65	Printing In Progress A		s_264	US bit	1 = True, 0 = False.  Read Only.

### Status of Scale B

09	Motion Out B			s_208	US bit	1 = True, 0 = False.  Read only.

0a	Center Of Zero B			s_209	US bit	1 = True, 0 = False.  Read only.

0b	Over Capacity B			s_20a	US bit	1 = True, 0 = False.  Read only.

0c	Under Zero B			s_20b	US bit	1 = True, 0 = False.  Read only.

0d	Net Mode B			s_20c	US bit	1 = True, 0 = False.  Read only.

0e	Scale Critical Error B		s_20d	US bit	1 = True, 0 = False.  Read only.

0f	Stored Weight Mode B		s_20e	US bit	1 = True, 0 = False.  Read only.

10	Scale Selected B			s_20f	US bit	1 = True, 0 = False.  Read only.

IDNET In Motion Error B		s_268 	US bit	Reserved for JagX O/S Use Only.

70	Weight Data OK B		s_269	US bit	1 = True, 0 = False.  Read Only.

6b	Rate Setpoint OK B		s_26a	US bit	1 = True, 0 = False.  Read Only.

6c	Estimated Weight B		s_26b	US bit	1 = True, 0 = False.  Read Only.

6d	Printing In Progress B		s_26c	US bit	1 = True, 0 = False.  Read Only.

### Status of Scale C

71	Motion Out C			s_270	US bit	1 = True, 0 = False.  Read only.

72	Center Of Zero C			s_271	US bit	1 = True, 0 = False.  Read only.

73	Over Capacity C			s_272	US bit	1 = True, 0 = False.  Read only.

74	Under Zero C			s_273	US bit	1 = True, 0 = False.  Read only.

75	Net Mode C			s_274	US bit	1 = True, 0 = False.  Read only.

76	Scale Critical Error C		s_275	US bit	1 = True, 0 = False.  Read only.

77	Stored Weight Mode C		s_276	US bit	1 = True, 0 = False.  Read only.

78	Scale Selected C			s_277	US bit	1 = True, 0 = False.  Read only.

IDNET In Motion Error C		s_278 	US bit	Reserved for JagX O/S Use Only.

7a	Weight Data OK C		s_279	US bit	1 = True, 0 = False.  Read Only.

7b	Rate Setpoint OK C		s_27a	US bit	1 = True, 0 = False.  Read Only.

7c	Estimated Weight C		s_27b	US bit	1 = True, 0 = False.  Read Only.

7d	Printing In Progress C		s_27c	US bit	1 = True, 0 = False.  Read Only.

### Status of Scale D

81	Motion Out D			s_280	US bit	1 = True, 0 = False.  Read only.

82	Center Of Zero D			s_281	US bit	1 = True, 0 = False.  Read only.

83	Over Capacity D			s_282	US bit	1 = True, 0 = False.  Read only.

84	Under Zero D			s_283	US bit	1 = True, 0 = False.  Read only.

85	Net Mode D			s_284	US bit	1 = True, 0 = False.  Read only.

86	Scale Critical Error D		s_285	US bit	1 = True, 0 = False.  Read only.

87	Stored Weight Mode D		s_286	US bit	1 = True, 0 = False.  Read only.

88	Scale Selected D			s_287	US bit	1 = True, 0 = False.  Read only.

IDNET In Motion Error D		s_288 	US bit	Reserved for JagX O/S Use Only.

8a	Weight Data OK D		s_289	US bit	1 = True, 0 = False.  Read Only.

8b	Rate Setpoint OK D		s_28a	US bit	1 = True, 0 = False.  Read Only.

8c	Estimated Weight D		s_28b	US bit	1 = True, 0 = False.  Read Only.

8d	Printing In Progress D		s_28c	US bit	1 = True, 0 = False.  Read Only.

### Status of Scale E (Summing Scale)

f1	Motion Out E			s_2f0	US bit	1 = True, 0 = False.  Read only.

f2	Center Of Zero E			s_2f1	US bit	1 = True, 0 = False.  Read only.

f3	Over Capacity E			s_2f2	US bit	1 = True, 0 = False.  Read only.

f4	Under Zero E			s_2f3	US bit	1 = True, 0 = False.  Read only.

f5	Net Mode E			s_2f4	US bit	1 = True, 0 = False.  Read only.

f6	Scale Critical Error E		s_2f5	US bit	1 = True, 0 = False.  Read only.

f7	Stored Weight Mode E		s_2f6	US bit	1 = True, 0 = False.  Read only.

f8	Scale Selected E			s_2f7	US bit	1 = True, 0 = False.  Read only.

IDNET In Motion Error E		s_2f8 	US bit	Reserved for JagX O/S Use Only.

fa	Weight Data OK E		s_2f9	US bit	1 = True, 0 = False.  Read Only.

fb	Rate Setpoint OK E		s_2fa	US bit	1 = True, 0 = False.  Read Only.

fc	Estimated Weight E		s_2fb	US bit	1 = True, 0 = False.  Read Only.

fd	Printing In Progress E		s_2fc	US bit	1 = True, 0 = False.  Read Only.

### Status of Setpoints 1-12

11	Setpoint Feeding 1		s_210	US bit	1 = True, 0 = False.  Read only.

12	Setpoint Fast Feeding 1		s_211	US bit	1 = True, 0 = False.  Read only.

13	Setpoint Within Tolerance 1	s_212	US bit	1 = True, 0 = False.  Read only.

15	Setpoint Feeding 2		s_214	US bit	1 = True, 0 = False.  Read only.

16	Setpoint Fast Feeding 2		s_215	US bit	1 = True, 0 = False.  Read only.

17	Setpoint Within Tolerance 2	s_216	US bit	1 = True, 0 = False.  Read only.

19	Setpoint Feeding 3		s_218	US bit	1 = True, 0 = False.  Read only.

1a	Setpoint Fast Feeding 3		s_219	US bit	1 = True, 0 = False.  Read only.

1b	Setpoint Within Tolerance 3	s_21a	US bit	1 = True, 0 = False.  Read only.

1d	Setpoint Feeding 4		s_21c	US bit	1 = True, 0 = False.  Read only.

1e	Setpoint Fast Feeding 4		s_21d	US bit	1 = True, 0 = False.  Read only.

1f	Setpoint Within Tolerance 4	s_21e	US bit	1 = True, 0 = False.  Read only.

21	Setpoint Feeding 5		s_220	US bit	1 = True, 0 = False.  Read only.

22	Setpoint Fast Feeding 5		s_221	US bit	1 = True, 0 = False.  Read only.

23	Setpoint Within Tolerance 5	s_222	US bit	1 = True, 0 = False.  Read only.

25	Setpoint Feeding6		s_224	US bit	1 = True, 0 = False.  Read only.

26	Setpoint Fast Feeding 6		s_225	US bit	1 = True, 0 = False.  Read only.

27	Setpoint Within Tolerance 6	s_226	US bit	1 = True, 0 = False.  Read only.

29	Setpoint Feeding 7		s_228	US bit	1 = True, 0 = False.  Read only.

2a	Setpoint Fast Feeding 7		s_229	US bit	1 = True, 0 = False.  Read only.

2b	Setpoint Within Tolerance 7	s_22a	US bit	1 = True, 0 = False.  Read only.

2d	Setpoint Feeding 8		s_22c	US bit	1 = True, 0 = False.  Read only.

2e	Setpoint Fast Feeding 8		s_22d	US bit	1 = True, 0 = False.  Read only.

2f	Setpoint Within Tolerance 8	s_22e	US bit	1 = True, 0 = False.  Read only.

31	Setpoint Feeding 9		s_230	US bit	1 = True, 0 = False.  Read only.

32	Setpoint Fast Feeding 9		s_231	US bit	1 = True, 0 = False.  Read only.

33	Setpoint Within Tolerance 9	s_232	US bit	1 = True, 0 = False.  Read only.

35	Setpoint Feeding 10		s_234	US bit	1 = True, 0 = False.  Read only.

36	Setpoint Fast Feeding 10		s_235	US bit	1 = True, 0 = False.  Read only.

37	Setpoint Within Tolerance 10	s_236	US bit	1 = True, 0 = False.  Read only.

39	Setpoint Feeding 11		s_238	US bit	1 = True, 0 = False.  Read only.

3a	Setpoint Fast Feeding 11		s_239	US bit	1 = True, 0 = False.  Read only.

3b	Setpoint Within Tolerance 11	s_23a	US bit	1 = True, 0 = False.  Read only.

3d	Setpoint Feeding 12		s_23c	US bit	1 = True, 0 = False.  Read only.

3e	Setpoint Fast Feeding 12		s_23d	US bit	1 = True, 0 = False.  Read only.

3f	Setpoint Within Tolerance 12	s_23e	US bit	1 = True, 0 = False.  Read only.

### Status of Remote Node Connections

Node OnLine_1			s_241	US bit 	Not used in Q.iMPACT.

Node OnLine_2			s_242	US bit 	Not used in Q.iMPACT.

Node OnLine_3			s_243	US bit 	Not used in Q.iMPACT.

Node OnLine_4			s_244	US bit 	Not used in Q.iMPACT.

Node OnLine_5			s_245	US bit 	Not used in Q.iMPACT.

Node OnLine_6			s_246	US bit 	Not used in Q.iMPACT.

PLC Online			s_24a	US bit 	Not used in Q.iMPACT.

4e	Host OnLine_3			s_24d	US bit 	Not used in Q.iMPACT.

4f	Host OnLine_2			s_24e	US bit 	Not used in Q.iMPACT.

50	Host OnLine_1			s_24f	US bit 	Not used in Q.iMPACT.

### JagBasic Applications Real-Time Status for a PLC or PC.

51	PLC_CustomStatus1_Scale A	s_250	US bit	1 = True, 0 = False.  Read/Write.

52	PLC_CustomStatus2_Scale A	s_251	US bit	1 = True, 0 = False.  Read/Write.

53	PLC_CustomStatus1_Scale B	s_252	US bit	1 = True, 0 = False.  Read/Write.

54	PLC_CustomStatus2_Scale B	s_253	US bit	1 = True, 0 = False.  Read/Write.

55	PLC_CustomStatus1_Scale C	s_254	US bit	1 = True, 0 = False.  Read/Write.

56	PLC_CustomStatus2_Scale C	s_255	US bit	1 = True, 0 = False.  Read/Write.

57	PLC_CustomStatus1_Scale D	s_256	US bit	1 = True, 0 = False.  Read/Write.

58	PLC_CustomStatus2_Scale D	s_257	US bit	1 = True, 0 = False.  Read/Write.

59	PLC_CustomStatus1_Scale E	s_258	US bit	1 = True, 0 = False.  Read/Write.

5a	PLC_CustomStatus2_Scale E	s_259	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25a	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25b	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25c	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25d	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25e	US bit	1 = True, 0 = False.  Read/Write.

JagBasic Custom Status 1		s_25f	US bit	1 = True, 0 = False.  Read/Write.

### Command Status Bits for Scale A

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

91	Tare Scale Error A		s_290	US bit	1 = True, 0 = False.  Read only.

92	Clear Tare Scale Error A		s_291	US bit	1 = True, 0 = False.  Read only.

93	Print Scale Error A		s_292	US bit	1 = True, 0 = False.  Read only.

94	Zero Scale Error A		s_293	US bit	1 = True, 0 = False.  Read only.

95	Switch to Prim Units Error A	s_294	US bit	1 = True, 0 = False.  Read only.

96	Switch to Second Units Error A	s_295	US bit	1 = True, 0 = False.  Read only.

97	Switch to Other Units Error A	s_296	US bit	1 = True, 0 = False.  Read only.

98	Apply Setup Error A		s_297	US bit	1 = True, 0 = False.  Read only.

99	Restart Setpoints Error A		s_298	US bit	1 = True, 0 = False.  Read only.

9a	Restart Rate Calculation Error A	s_299	US bit	1 = True, 0 = False.  Read only.

9b	Restart Filter Error A		s_29a	US bit	1 = True, 0 = False.  Read only.

9c	Reset Setpoint Coincidence Error A	s_29b	US bit	1 = True, 0 = False.  Read only.

9d	Disable Scale Error A		s_29c	US bit	1 = True, 0 = False.  Read only.

9e	Capture Raw Counts Error A	s_29d	US bit	1 = True, 0 = False.  Read only.

9f	Write Cal. To EEProm Error A	s_29e	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for Scale B

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

a1	Tare Scale Error B		s_2a0	US bit	1 = True, 0 = False.  Read only.

a2	Clear Tare Scale Error B		s_2a1	US bit	1 = True, 0 = False.  Read only.

a3	Print Scale Error B		s_2a2	US bit	1 = True, 0 = False.  Read only.

a4	Zero Scale Error B		s_2a3	US bit	1 = True, 0 = False.  Read only.

a5	Switch to Prim Units Error B	s_2a4	US bit	1 = True, 0 = False.  Read only.

a6	Switch to Second Units Error B	s_2a5	US bit	1 = True, 0 = False.  Read only.

a7	Switch to Other Units Error B	s_2a6	US bit	1 = True, 0 = False.  Read only.

a8	Apply Setup Error B		s_2a7	US bit	1 = True, 0 = False.  Read only.

a9	Restart Setpoint Error B		s_2a8	US bit	1 = True, 0 = False.  Read only.

aa	Restart Rate Calculation Error B	s_2a9	US bit	1 = True, 0 = False.  Read only.

ab	Restart Filter Error B		s_2aa	US bit	1 = True, 0 = False.  Read only.

ac	Reset Setpoint Coincidence Error B	s_2ab	US bit	1 = True, 0 = False.  Read only.

ad	Disable Scale Error B		s_2ac	US bit	1 = True, 0 = False.  Read only.

ae	Capture Raw Counts Error B	s_2ad	US bit	1 = True, 0 = False.  Read only.

af	Write Cal. To EEProm Error B	s_2ae	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for the Selected Scale

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

b1	Tare Scale Error S		s_2b0	US bit	1 = True, 0 = False.  Read only.

b2	Clear Tare Scale Error S		s_2b1	US bit	1 = True, 0 = False.  Read only.

b3	Print Scale Error S		s_2b2	US bit	1 = True, 0 = False.  Read only.

b4	Zero Scale Error S		s_2b3	US bit	1 = True, 0 = False.  Read only.

b5	Switch to Prim Units Error S	s_2b4	US bit	1 = True, 0 = False.  Read only.

b6	Switch to Second Units Error S	s_2b5	US bit	1 = True, 0 = False.  Read only.

b7	Switch to Other Units Error S	s_2b6	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for Custom Print

b8	Custom Print Error_1		s_2b7	US bit	1 = True, 0 = False.  Read only.

b9	Custom Print Error_2		s_2b8	US bit	1 = True, 0 = False.  Read only.

ba	Custom Print Error_3		s_2b9	US bit	1 = True, 0 = False.  Read only.

bb	Custom Print Error_4		s_2ba	US bit	1 = True, 0 = False.  Read only.

bc	Custom Print Error_5		s_2bb	US bit	1 = True, 0 = False.  Read only.

Alarm Output			s_2be	US bit	1 = True, 0 = False.  Read only.

JagBasic Enabled		s_2bf	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for Scale C

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

c1	Tare Scale Error C		s_2c0	US bit	1 = True, 0 = False.  Read only.

c2	Clear Tare Scale Error C		s_2c1	US bit	1 = True, 0 = False.  Read only.

c3	Print Scale Error C		s_2c2	US bit	1 = True, 0 = False.  Read only.

c4	Zero Scale Error C		s_2c3	US bit	1 = True, 0 = False.  Read only.

c5	Switch to Prim Units Error C	s_2c4	US bit	1 = True, 0 = False.  Read only.

c6	Switch to Second Units Error C	s_2c5	US bit	1 = True, 0 = False.  Read only.

c7	Switch to Other Units Error C	s_2c6	US bit	1 = True, 0 = False.  Read only.

c8	Apply Setup Error C		s_2c7	US bit	1 = True, 0 = False.  Read only.

c9	Restart Setpoints Error C		s_2c8	US bit	1 = True, 0 = False.  Read only.

ca	Restart Rate Calculation Error C	s_2c9	US bit	1 = True, 0 = False.  Read only.

cb	Restart Filter Error C		s_2ca	US bit	1 = True, 0 = False.  Read only.

cc	Reset Setpoint Coincidence Error C	s_2cb	US bit	1 = True, 0 = False.  Read only.

cd	Disable Scale Error C		s_2cc	US bit	1 = True, 0 = False.  Read only.

ce	Capture Raw Counts Error C	s_2cd	US bit	1 = True, 0 = False.  Read only.

cf	Write Cal. To EEProm Error C	s_2ce	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for Scale D

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

d1	Tare Scale Error D		s_2d0	US bit	1 = True, 0 = False.  Read only.

d2	Clear Tare Scale Error D		s_2d1	US bit	1 = True, 0 = False.  Read only.

d3	Print Scale Error D		s_2d2	US bit	1 = True, 0 = False.  Read only.

d4	Zero Scale Error D		s_2d3	US bit	1 = True, 0 = False.  Read only.

d5	Switch to Prim Units Error D	s_2d4	US bit	1 = True, 0 = False.  Read only.

d6	Switch to Second Units Error D	s_2d5	US bit	1 = True, 0 = False.  Read only.

d7	Switch to Other Units Error D	s_2d6	US bit	1 = True, 0 = False.  Read only.

d8	Apply Setup Error D		s_2d7	US bit	1 = True, 0 = False.  Read only.

d9	Restart Setpoints Error D		s_2d8	US bit	1 = True, 0 = False.  Read only.

da	Restart Rate Calculation Error D	s_2d9	US bit	1 = True, 0 = False.  Read only.

db	Restart Filter Error D		s_2da	US bit	1 = True, 0 = False.  Read only.

dc	Reset Setpoint CoincidenceError D	s_2db	US bit	1 = True, 0 = False.  Read only.

dd	Disable Scale Error D		s_2dc	US bit	1 = True, 0 = False.  Read only.

de	Capture Raw Counts Error D	s_2dd	US bit	1 = True, 0 = False.  Read only.

df	Write Cal. To EEProm Error D	s_2de	US bit	1 = True, 0 = False.  Read only.

### Command Status Bits for Scale E

The JagX Operating System sets the following fields to report success (=0) or error (=1) when an application uses a corresponding discrete field to trigger a command in the JagX Operating System.

e1	Tare Scale Error E		s_2e0	US bit	1 = True, 0 = False.  Read only.

e2	Clear Tare Scale Error E		s_2e1	US bit	1 = True, 0 = False.  Read only.

e3	Print Scale Error E		s_2e2	US bit	1 = True, 0 = False.  Read only.

e4	Zero Scale Error E		s_2e3	US bit	1 = True, 0 = False.  Read only.

e5	Switch to Prim Units Error E	s_2e4	US bit	1 = True, 0 = False.  Read only.

e6	Switch to Second Units Error E	s_2e5	US bit	1 = True, 0 = False.  Read only.

e7	Switch to Other Units Error E	s_2e6	US bit	1 = True, 0 = False.  Read only.

e8	Apply Setup Error E		s_2e7	US bit	1 = True, 0 = False.  Read only.

e9	Restart Setpoints Error E		s_2e8	US bit	1 = True, 0 = False.  Read only.

ea	Restart Rate Calculation Error E	s_2e9	US bit	1 = True, 0 = False.  Read only.

eb	Restart Filter Error E		s_2ea	US bit	1 = True, 0 = False.  Read only.

ec	Reset Setpoint CoincidenceError E	s_2eb	US bit	1 = True, 0 = False.  Read only.

ed	Disable Scale Error E		s_2ec	US bit	1 = True, 0 = False.  Read only.

### Miscellaneous JagX O/S Triggers

Offset 0x600

01	Master Control Relay		t_600	US bit	1 = set physical discrete outputs active,

0 = set all physical discrete outputs to 0. Read/Write.

In-Setup Mode			s_601	US bit	Reserved for JagX O/S Use Only.

Disable PLC			t_602	US bit	Reserved for JagX O/S Use Only.

03	Disable Error Display		t_603	US bit	1 = True, 0 = False.  Read/Write.

04	Disable Numeric Display		t_604	US bit	1 = True, 0 = False.  Read/Write.

AnalogOut Trim Adjust		t_605	US bit	Reserved for JagX O/S Use Only.

06	Restart Network			s_606	US bit	1 = True, 0 = False.  Read/Write.

AnalogOut Use Zero Preset	s_607	US bit	Reserved for JagX O/S Use Only.

AnalogOut Use Trim Preset	s_608	US bit 	Reserved for JagX O/S Use Only.

0a	Disable Setup			t_609	US bit	1 = True, 0 = False.  Read/Write.

0b	Disable Keypad			t_60a	US bit	1 = True, 0 = False.  Read/Write.

0c	Ignore QWERTY Position Keys	t_60b	US bit	1 = True, 0 = False.  Read/Write.

Internal Disable Error Display	t_60c	US bit 	Reserved for JagX O/S Use Only.

0e	Disable RunFlat			t_60d	US bit 	1 = True, 0 = False.  Read/Write.

0f	Alarm Output Acknowledge	t_60e	US bit 	1 = True, 0 = False.  Read/Write.

Maintenance Only Mode In Setup	s_60f	US bit 	1 = True, 0 = False.  Read/Write.

## Edge Sensitive Discrete Status Object

ControlNet Class Code: 76 hex

- The Edge Sensitive Discrete Status Object has only a single instance, instance 1.

- NOTE: This object’s attributes are offset by 600hex (0x600); Actual Data = (0x600 + Attr. No.)

EDGE SENSITIVE LOGICAL DISCRETE I/O DATA	(External Read/Write)

Edge sensitive bit fields only trigger events when a 1 is written to the field.

### Miscellaneous JagX O/S Internal Triggers

Attr. 	Function/Value			Name	Format	Description

Run Scanner			t_680	US bit	Reserved for JagX O/S Use Only.

Stop Scanner			t_681	US bit	Reserved for JagX O/S Use Only.

Select Key			t_682	US bit	Reserved for JagX O/S Use Only.

Escape Key			t_683	US bit	Reserved for JagX O/S Use Only.

Enter Key			t_684	US bit	Reserved for JagX O/S Use Only.

Alarm				t_685	US bit	Reserved for JagX O/S Use Only.

Update Display A			t_686	US bit	Reserved for JagX O/S Use Only.

Update Display B			t_687	US bit	Reserved for JagX O/S Use Only.

Update Display C			t_610	US bit	Reserved for JagX O/S Use Only.

Update Display D			t_611	US bit	Reserved for JagX O/S Use Only.

Update Display E			t_612	US bit	Reserved for JagX O/S Use Only.

### JagX O/S Event Indications

The JagX O/S sets these discrete triggers to indicate an event.  A JagBasic application can use JagBasic events to monitor these O/S events .  The JagBasic application must set the field to 0 before the same event will trigger again.

88	Weight Updated A		t_688	US bit	1 = Trigger, 0 = Complete.  Read/Write.

89	Weight Updated B		t_689	US bit	1 = Trigger, 0 = Complete.  Read/Write.

8a	Flow Meter Updated		t_68a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

8b	Web Pages Disable Error Display	t_68b	US bit	1=WebPages Disables Error Display

8c	Setpoint Installed A		t_68c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

8d	Setpoint Installed B		t_68d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

8e	Calibration Complete A		t_68e	US bit	1 = Trigger, 0 = Complete.  Read/Write.

8f	Calibration Complete B		t_68f	US bit	1 = Trigger, 0 = Complete.  Read/Write.

13	Weight Updated C		t_613	US bit	1 = Trigger, 0 = Complete.  Read/Write.

14	Weight Updated D		t_614	US bit	1 = Trigger, 0 = Complete.  Read/Write.

15	Weight Updated E		t_615	US bit	1 = Trigger, 0 = Complete.  Read/Write.

16	Setpoint Installed C		t_616	US bit	1 = Trigger, 0 = Complete.  Read/Write.

17	Setpoint Installed D		t_617	US bit	1 = Trigger, 0 = Complete.  Read/Write.

18	Setpoint Installed E		t_618	US bit	1 = Trigger, 0 = Complete.  Read/Write.

19	Calibration Complete C		t_619	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1a	Calibration Complete D		t_61a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1b	Calibration Complete E		t_61b	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1c	Start Setpoint Jog Timer		t_61c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1d	Start Prompt List			t_61d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1e	Request to Stop JagBasic Pgm	t_61e	US bit	1 = Trigger, 0 = Complete.  Read/Write.

1f	Setup Started by Web Browser	t_61f	US bit	Reserved for Q.iMPACT O/S Only.

### Command Triggers for Scale A

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

90	Tare Scale A			t_690	US bit	1 = Trigger, 0 = Complete.  Read/Write.

91	Clear Tare Scale A		t_691	US bit	1 = Trigger, 0 = Complete.  Read/Write.

92	Print Scale A			t_692	US bit	1 = Trigger, 0 = Complete.  Read/Write.

93	Zero Scale A			t_693	US bit	1 = Trigger, 0 = Complete.  Read/Write.

94	Switch to Primary Units A		t_694	US bit	1 = Trigger, 0 = Complete.  Read/Write.

95	Switch to Second Units A		t_695	US bit	1 = Trigger, 0 = Complete.  Read/Write.

96	Switch to Other Units A		t_696	US bit	1 = Trigger, 0 = Complete.  Read/Write.

97	Apply Setup A			t_697	US bit	1 = Trigger, 0 = Complete.  Read/Write.

98	Restart Setpoints A		t_698	US bit	1 = Trigger, 0 = Complete.  Read/Write.

99	Restart Rate Calculation A		t_699	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9a	Restart Filter A			t_69a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9b	Reset Setpoint Coincidence A	t_69b	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9c	Disable Scale A			t_69c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9d	Capture Raw Counts A		t_69d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9e	Write Calibration to EEProm A	t_69e	US bit	1 = Trigger, 0 = Complete.  Read/Write.

9f	Reset Predictive Failures		t_69f	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Command Triggers for Scale B

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

a0	Tare Scale B			t_6a0	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a1	Clear Tare Scale B		t_6a1	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a2	PrintScale B			t_6a2	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a3	Zero Scale B			t_6a3	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a4	Switch to Primary Units B		t_6a4	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a5	Switch to Second Units B		t_6a5	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a6	Switch to Other Units B		t_6a6	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a7	Apply Setup B			t_6a7	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a8	Restart Setpoints B		t_6a8	US bit	1 = Trigger, 0 = Complete.  Read/Write.

a9	Restart Rate Calculation B		t_6a9	US bit	1 = Trigger, 0 = Complete.  Read/Write.

aa	Restart Filter B			t_6aa	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ab	Reset Setpoint Coincidence B	t_6ab	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ac	Disable Scale B			t_6ac	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ad	Capture Raw Counts B		t_6ad	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ae	Write Calibration to EEProm B	t_6ae	US bit	1 = Trigger, 0 = Complete.  Read/Write.

af	Reset Predictive Failures		t_6af	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Command Triggers for the Selected Scale

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

b0	Tare Scale S			t_6b0	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b1	Clear Tare Scale S		t_6b1	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b2	Print Scale S			t_6b2	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b3	Zero Scale S			t_6b3	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b4	Switch to Primary Units S		t_6b4	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b5	Switch to Second Units S		t_6b5	US bit	1 = Trigger, 0 = Complete.  Read/Write.

b6	Switch to Other Units S		t_6b6	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Command Triggers for the Scale C

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

20	Tare Scale C			t_620	US bit	1 = Trigger, 0 = Complete.  Read/Write.

21	Clear Tare Scale C		t_621	US bit	1 = Trigger, 0 = Complete.  Read/Write.

22	Print Scale C			t_622	US bit	1 = Trigger, 0 = Complete.  Read/Write.

23	Zero Scale C			t_623	US bit	1 = Trigger, 0 = Complete.  Read/Write.

24	Switch to Primary Units C		t_624	US bit	1 = Trigger, 0 = Complete.  Read/Write.

25	Switch to Second Units C		t_625	US bit	1 = Trigger, 0 = Complete.  Read/Write.

26	Switch to Other Units C		t_626	US bit	1 = Trigger, 0 = Complete.  Read/Write.

27	Apply Setup C			t_627	US bit	1 = Trigger, 0 = Complete.  Read/Write.

28	Restart Setpoints C		t_628	US bit	1 = Trigger, 0 = Complete.  Read/Write.

29	Restart Rate Calculation C		t_629	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2a	Restart Filter C			t_62a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2b	Reset Setpoint Coincidence C	t_62b	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2c	Disable Scale C			t_62c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2d	Capture Raw Counts C		t_62d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2e	Write Calibration to EEProm C	t_62e	US bit	1 = Trigger, 0 = Complete.  Read/Write.

2f	Reset Predictive Failures		t_62f	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Command Triggers for the Scale D

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

30	Tare Scale D			t_630	US bit	1 = Trigger, 0 = Complete.  Read/Write.

31	Clear Tare Scale D		t_631	US bit	1 = Trigger, 0 = Complete.  Read/Write.

32	Print Scale D			t_632	US bit	1 = Trigger, 0 = Complete.  Read/Write.

33	Zero Scale D			t_633	US bit	1 = Trigger, 0 = Complete.  Read/Write.

34	Switch to Primary Units D		t_634	US bit	1 = Trigger, 0 = Complete.  Read/Write.

35	Switch to Second Units D		t_635	US bit	1 = Trigger, 0 = Complete.  Read/Write.

36	Switch to Other Units D		t_636	US bit	1 = Trigger, 0 = Complete.  Read/Write.

37	Apply Setup D			t_637	US bit	1 = Trigger, 0 = Complete.  Read/Write.

38	Restart Setpoints D		t_638	US bit	1 = Trigger, 0 = Complete.  Read/Write.

39	Restart Rate Calculation D		t_639	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3a	Restart Filter D			t_63a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3b	Reset Setpoint Coincidence D	t_63b	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3c	Disable Scale D			t_63c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3d	Capture Raw Counts D		t_63d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3e	Write Calibration to EEProm D	t_63e	US bit	1 = Trigger, 0 = Complete.  Read/Write.

3f	Reset Predictive Failures		t_63f	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Command Triggers for the Scale E  (Summing Scale)

Applications can set the following fields to trigger a command within the JagX O/S.  The JagX O/S will set the field to 0 when it is done processing the command, and it will set a corresponding error bit to indicate when there is an error in processing the command.

40	Tare Scale E			t_640	US bit	1 = Trigger, 0 = Complete.  Read/Write.

41	Clear Tare Scale E		t_641	US bit	1 = Trigger, 0 = Complete.  Read/Write.

42	Print Scale E			t_642	US bit	1 = Trigger, 0 = Complete.  Read/Write.

43	Zero Scale E			t_643	US bit	1 = Trigger, 0 = Complete.  Read/Write.

44	Switch to Primary Units E		t_644	US bit	1 = Trigger, 0 = Complete.  Read/Write.

45	Switch to Second Units E		t_645	US bit	1 = Trigger, 0 = Complete.  Read/Write.

46	Switch to Other Units E		t_646	US bit	1 = Trigger, 0 = Complete.  Read/Write.

47	Apply Setup E			t_647	US bit	1 = Trigger, 0 = Complete.  Read/Write.

48	Restart Setpoints E		t_648	US bit	1 = Trigger, 0 = Complete.  Read/Write.

49	Restart Rate Calculation E		t_649	US bit	1 = Trigger, 0 = Complete.  Read/Write.

4a	Restart Filter E			t_64a	US bit	1 = Trigger, 0 = Complete.  Read/Write.

4b	Reset Setpoint Coincidence E	t_64b	US bit	1 = Trigger, 0 = Complete.  Read/Write.

4c	Disable Scale E			t_64c	US bit	1 = Trigger, 0 = Complete.  Read/Write.

4d	Capture Raw Counts E		t_64d	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Select Scale Command Triggers

c0	Select Scale A			t_6c0	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c1	Select Scale B			t_6c1	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c2	Select Other Scale		t_6c2	US bit	1 = Trigger, 0 = Complete.  Read/Write.

50	Select Scale C			t_650	US bit	1 = Trigger, 0 = Complete.  Read/Write.

51	Select Scale D			t_651	US bit	1 = Trigger, 0 = Complete.  Read/Write.

52	Select Scale E			t_652	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Custom Print Command Triggers

c3	Demand Custom Print_1		t_6c3	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c4	Demand Custom Print_2		t_6c4	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c5	Demand Custom Print_3		t_6c5	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c6	Demand Custom Print_4		t_6c6	US bit	1 = Trigger, 0 = Complete.  Read/Write.

c7	Demand Custom Print_5		t_6c7	US bit	1 = Trigger, 0 = Complete.  Read/Write.

Continuous Custom Print A	t_6c8	US bit	Reserved for JagX O/S Use Only.

Continuous Custom Print B	t_6c9	US bit	Reserved for JagX O/S Use Only.

Continuous Custom Print C	t_653	US bit	Reserved for JagX O/S Use Only.

Continuous Custom Print D	t_654	US bit	Reserved for JagX O/S Use Only.

Continuous Custom Print E	t_655	US bit	Reserved for JagX O/S Use Only.

### Miscellaneous Command Triggers

56	EmailAlertMessage_1		t_656	US bit	1 = Trigger, 0 = Complete.  Read/Write.

57	EmailAlertMessage_2		t_657	US bit	1 = Trigger, 0 = Complete.  Read/Write.

58	EmailAlertMessage_3		t_658	US bit	1 = Trigger, 0 = Complete.  Read/Write.

59	EmailAlertMessage_4		t_659	US bit	1 = Trigger, 0 = Complete.  Read/Write.

Reset Serial Port Parameters_1	t_65c	US bit 	Reserved for JagX O/S Use Only.

Reset Serial Port Parameters_2	t_65d	US bit 	Reserved for JagX O/S Use Only.

Reset Serial Port Parameters_3	t_65e	US bit 	Reserved for JagX O/S Use Only.

Reset Serial Port Parameters_4	t_65f	US bit 	Reserved for JagX O/S Use Only.

### PC Host Real-Time Commands for a JagBasic Application

cc	Custom Command 1		t_6cc	US bit	1 = Trigger, 0 = Complete.  Read/Write.

cd	Custom Command 2		t_6cd	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ce	Custom Command 3		t_6ce	US bit	1 = Trigger, 0 = Complete.  Read/Write.

cf	Custom Command 4		t_6cf	US bit	1 = Trigger, 0 = Complete.  Read/Write.

### Reserved  Triggers for JagX O/S

Basic Program Break		t_6d0	US bit 	Reserved for JagX O/S Use Only.

Start Basic Program_1		t_6d1	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_2		t_6d2	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_3		t_6d3	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_4		t_6d4	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_5		t_6d5	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_6		t_6d6	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_7		t_6d7	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_8		t_6d8	US bit	Reserved for JagX O/S Use Only.

Start Basic Program_9		t_6d9	US bit	Reserved for JagX O/S Use Only.

Basic Execution Running		t_6da	US bit	Reserved for JagX O/S Use Only.

Basic Awaiting Key Input		t_6db	US bit	Reserved for JagX O/S Use Only.

Basic Keystroke Ready		t_6dc	US bit	Reserved for JagX O/S Use Only.

Send BRAM Complete		t_6dd	US bit	Reserved for JagX O/S Use Only.

Reset Serial Connections		t_6de	US bit	Reserved for JagX O/S Use Only.

Prompt List Running		t_6df	US bit	Reserved for JagX O/S Use Only.

## Rising/Falling Input Edge-Sensitive Object

ControlNet Class Code: 77 hex

The Edge Sensitive Physical Discrete Input Object has only a single instance, instance 1.

- NOTE: This object’s attributes  are offset by 600hex (0x600); Actual Data = (0x600 + Attr. No.)

The JagX O/S sets these fields=1 whenever it detects a rising or falling edge in the discrete inputs.  A JagBasic application can use events to monitor these fields.  It must set the field to 0 before the same event will trigger again.

Attr. 	Function/Value			Name	Format	Description

e0	Discrete Input Rising Edge_1	p_6e0	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e1	Discrete Input Rising Edge_2	p_6e1	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e2	Discrete Input Rising Edge_3	p_6e2	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e3	Discrete Input Rising Edge_4	p_6e3	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e4	Discrete Input Rising Edge_5	p_6e4	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e5	Discrete Input Rising Edge_6	p_6e5	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e6	Discrete Input Rising Edge_7	p_6e6	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e7	Discrete Input Rising Edge_8	p_6e7	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e8	Discrete Input Rising Edge_9	p_6e8	US bit	1 = Trigger, 0 = Complete.  Read/Write.

e9	Discrete Input Rising Edge_10	p_6e9	US bit	1 = Trigger, 0 = Complete.  Read/Write.

ea	Discrete Input Rising Edge_11	p_6ea	US bit	1 = Trigger, 0 = Complete.  Read/Write.

eb	Discrete Input Rising Edge_12	p_6eb	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f0	Discrete Input Falling Edge_1	p_6f0	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f1	Discrete Input Falling Edge_2	p_6f1	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f2	Discrete Input Falling Edge_3	p_6f2	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f3	Discrete Input Falling Edge_4	p_6f3	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f4	Discrete Input Falling Edge_5	p_6f4	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f5	Discrete Input Falling Edge_6	p_6f5	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f6	Discrete Input Falling Edge_7	p_6f6	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f7	Discrete Input Falling Edge_8	p_6f7	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f8	Discrete Input Falling Edge_9	p_6f8	US bit	1 = Trigger, 0 = Complete.  Read/Write.

f9	Discrete Input Falling Edge_10	p_6f9	US bit	1 = Trigger, 0 = Complete.  Read/Write.

fa	Discrete Input Falling Edge_11	p_6fa	US bit	1 = Trigger, 0 = Complete.  Read/Write.

fb	Discrete Input Falling Edge_12	p_6fb	US bit	1 = Trigger, 0 = Complete.  Read/Write.

## Level Sensitive Physical Discrete I/O Object

ControlNet Class Code: 78 hex

- The Level Sensitive Physical Discrete I/O Object has only a single instance, instance 1.

- NOTE: The input  attributes  are offset by 100hex (0x100); Actual Data = (0x100 + (Attr. No. –1))

- NOTE: The output  attributes  are offset by 500hex (0x500); Actual Data = (0x500 + (Attr. No.-1))

Attr. 	Function/Value			Name	Format	Description

PHYSICAL DISCRETE I/O DATA

01	Physical Discrete Input_1		p_100	US bit	1 = True, 0 = False.  Read Only.  Controller.

02	Physical Discrete Input_2		p_101	US bit	1 = True, 0 = False.  Read Only.  Controller.

03	Physical Discrete Input_3		p_102	US bit	1 = True, 0 = False.  Read Only.  Controller.

04	Physical Discrete Input_4		p_103	US bit	1 = True, 0 = False.  Read Only.  Controller.

09	Physical Discrete Input_5		p_108	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0a	Physical Discrete Input_6		p_109	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0b	Physical Discrete Input_7 		p_10a	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0c	Physical Discrete Input_8 		p_10b	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0d	Physical Discrete Input_9 		p_10c	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0e	Physical Discrete Input_10		p_10d	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

0f	Physical Discrete Input_11		p_10e	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

10	Physical Discrete Input_12		p_10f	US bit	1 = True, 0 = False.  Read Only.  MFIO #1.

- 01	Physical Discrete Output_1	p_500	US bit	1 = True, 0 = False.  Read/Write.  Controller. 02	Physical Discrete Output_2	p_501	US bit	1 = True, 0 = False.  Read/Write.  Controller.

03	Physical Discrete Output_3	p_502	US bit	1 = True, 0 = False.  Read/Write.  Controller.

04	Physical Discrete Output_4	p_503	US bit	1 = True, 0 = False.  Read/Write.  Controller.

09	Physical Discrete Output_5	p_508	US bit	1 = True, 0 = False.  Read/Write.  MFIO.

0a	Physical Discrete Output_6	p_509	US bit	1 = True, 0 = False.  Read/Write.  Outputs

0b	Physical Discrete Output_7	p_50a	US bit	1 = True, 0 = False.  Read/Write.  start

0c	Physical Discrete Output_8	p_50b	US bit	1 = True, 0 = False.  Read/Write.  after MFIO

0d	Physical Discrete Output_9 	p_50c	US bit	1 = True, 0 = False.  Read/Write.  Inputs.

0e	Physical Discrete Output_10 	p_50d	US bit	1 = True, 0 = False.  Read/Write. “jag13”

0f	Physical Discrete Output_11	p_50e	US bit	1 = True, 0 = False.  Read/Write. defines the

10	Physical Discrete Output_12	p_50f	US bit	1 = True, 0 = False.  Read/Write. number

Discrete Inputs.

# SPECIAL REMOTE ACCESS FILE NAMES

_vfd1	Read contents of JagX display.  The returned JagX Message data field is 40 bytes long.  The first 18 bytes have the contents of the numeric display.  The second 18 bytes have the contents of

the alphanumeric display.  The next two bytes have the contents of the numeric cursors.  The final two bytes have the contents of the alphanumeric cursors.

_frst	Read the first shared data item from Heap, BRAM, and EEPROM.

_next	Read the next shared data item from Heap, BRAM, and EEPROM

_list	Read a list of Shared Data items.  The JagX Message request contains a list of Shared Data

items.  The returned JagX Message data field contains an ordered list of the contents of the

Shared Data items.  Each response item contains the data type, length, and contents of each

field.

_din1	Read all discrete inputs with one remote access.  The returned JagX Message data field is a

2-byte field with each bit representing one discrete input. This field cannot be read through  Web

Server --- use bd030.

_o1ds	Read all discrete outputs with one remote access.  The returned JagX Message data field is a

2-byte field with each bit representing one discrete output. This field cannot be read through  Web

Server --- use bd031.

_blnn	Read contents of a block of discrete global data.  “nn” is the table number of the starting discrete.

The returned JagX Message data field is 2-bytes long.

# Q.iMPACT Shared Data

- The Q.iMPACT System Database defines the Q.iMPACT system configuration, the PAC parameters for each Measuring Device and Material-Path, and the processing commands and statuses.

## Global Cluster Configuration

### Q.iMPACT Cluster Terminals Table

There can be up to 20 Q.iMPACT Network Terminals in an PAC System.  The IP address for each of these terminal must be specifically set up for each terminal that is present in the cluster.

There can be up to 3 specific Hosts.  If there are no specific hosts set up in the Network Terminals table, ANY Host client may connect to Q.iMPACT Shared Data Servers.

ControlNet Class Code: 80 hex

- There is one instance of the Q.iMPACT Cluster Terminals Table.

Attr. 	Function/Value			Name	Format	Description

11	Cluster Terminal IP Address _ 1	nde11	UL	Setup.  Default is 111.111.111.111.

12	Cluster Terminal IP Address _ 2	nde12	UL	Setup.  Default is 111.111.111.112.

13	Cluster Terminal IP Address _ 3	nde13	UL	Setup.  Default is 111.111.111.113.

14	Cluster Terminal IP Address _ 4	nde14	UL	Setup.  Default is 111.111.111.114.

15	Cluster Terminal IP Address _ 5	nde15	UL	Setup.  Default is 111.111.111.115.

16	Cluster Terminal IP Address _ 6	nde16	UL	Setup.  Default is 111.111.111.116.

17	Cluster Terminal IP Address _ 7	nde17	UL	Setup.  Default is 111.111.111.117.

18	Cluster Terminal IP Address _ 8	nde18	UL	Setup.  Default is 111.111.111.118.

19	Cluster Terminal IP Address _ 9	nde19	UL	Setup.  Default is 111.111.111.119.

20	Cluster Terminal IP Address _ 10	nde20	UL	Setup.  Default is 111.111.111.120.

21	Cluster Terminal IP Address _ 11	nde21	UL	Setup.  Default is 111.111.111.121.

22	Cluster Terminal IP Address _ 12	nde22	UL	Setup.  Default is 111.111.111.122.

23	Cluster Terminal IP Address _ 13	nde23	UL	Setup.  Default is 	111.111.111.123.

24	Cluster Terminal IP Address _ 14	nde24	UL	Setup.  Default is 	111.111.111.124.

25	Cluster Terminal IP Address _ 15	nde25	UL	Setup.  Default is 	111.111.111.125.

26	Cluster Terminal IP Address _ 16	nde26	UL	Setup.  Default is 	111.111.111.126.

27	Cluster Terminal IP Address _ 17	nde27	UL	Setup.  Default is 111.111.111.127.

28	Cluster Terminal IP Address _ 18	nde28	UL	Setup.  Default is 	111.111.111.128.

29	Cluster Terminal IP Address _ 19	nde29	UL	Setup.  Default is 	111.111.111.129.

30	Cluster Terminal IP Address _ 20	nde30	UL	Setup.  Default is 	111.111.111.130.

31	Host IP Address _ 1		nde31	UL	Setup.  Default is 	000.000.000.000.

32	Host IP Address _ 2		nde32	UL	Setup.  Default is 	000.000.000.000.

33	Host IP Address _ 3		nde33	UL	Setup.  Default is 000.000.000.000.

### Q.iMPACT Terminal Enable Table

ControlNet Class Code: 81 hex

- There is one instance of the Q.iMPACT Terminal Enable Table.

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	Cluster Connection 1		rm101	UC	Setup.  0 = Disabled, 1 = Enabled

2	Cluster Connection 2		rm102	UC	Setup.  0 = Disabled, 1 = Enabled

3	Cluster Connection 3		rm103	UC	Setup.  0 = Disabled, 1 = Enabled

.			   .	 .		.

.			   .	 . 		.

19	Cluster Connection 19		rm119	UC	Setup.  0 = Disabled, 1 = Enabled

20	Cluster Connection 20		rm120	UC	Setup.  0 = Disabled, 1 = Enabled

21	Host Connection 1		rm121	UC	Setup.  0 = Disabled, 1 = Enabled

22	Host Connection 2		rm122	UC	Setup.  0 = Disabled, 1 = Enabled

23	Host Connection 3		rm123	UC	Setup.  0 = Disabled, 1 = Enabled

24	Reserved			rm124	UC	Setup.  0 = Disabled, 1 = Enabled

### Q.iMPACT Terminal Online Table

ControlNet Class Code: 6F hex

- There is one instance of the Q.iMPACT Terminal Enable Table.

(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	Cluster Connection 1		ro101	UC	Calculated. 0=Offline, 1=Online

2	Cluster Connection 2		ro102	UC	Calculated. 0=Offline, 1=Online

3	Cluster Connection 3		ro103	UC	Calculated. 0=Offline, 1=Online

.			   .	 .		.

.			   .	 . 		.

19	Cluster Connection 19		ro119	UC	Calculated. 0=Offline, 1=Online

20	Cluster Connection 20		ro120	UC	Calculated. 0=Offline, 1=Online

21	Host Connection 1		ro121	UC	Calculated. 0=Offline, 1=Online

22	Host Connection 2		ro122	UC	Calculated. 0=Offline, 1=Online

23	Host Connection 3		ro123	UC	Calculated. 0=Offline, 1=Online

24	PLC Connection			ro124	UC	Calculated. 0=Offline, 1=Online

### Q.iMPACT Instrument Channel Configuration

ControlNet Class Code: 82 hex

- There are 10 fields in the Q.iMPACT Channel Configuration.  There are up to 200 instrument channels in the cluster.  Each field contains the configuration data for up to 20 instrument channels in the cluster.  The first field represents channels 1-20.  The second field represents channels 21-40, etc.

- The instrument configuration contains 4 characters for each instrument  – its Terminal Identifier, its Process Identifier, its Unit number, and the Q.iMPACT Terminal containing its Controller Interface.

- PLC Measuring Device Channel Configuration SHARED DATA

Attr. 	Function/Value			Name	Format	Description

- 1	Instrument Channel String		ins01	80C	Setup. The string contains the channel of

- instruments 1-20 in the cluster.

- Each the instrument channel contains 4

- characters, as follows:

- Q.iMPACT terminal identifier.  @, 1-9, A-K represent Terminals None, 1-20.

- Instrument identifier.   @, A-D represents scales, K-V represents flow meters

- Monitor channel, @ = no monitor, M  = monitor, used by maintenance-channel web pages.

- Bridge Terminal that contains the Controller Interface Card for this instrument.  @, 1-9, A-K represent  Terminals None,  1-20.

- 2	Instrument Channel String		ins02	80C	Setup. The string contains the channel of

- instruments 21-40.

- 3	Instrument Channel String		ins03	80C	Setup. The string contains the channel of

- instruments 41-60.

- 4	Instrument Channel String		ins04	80C	Setup. The string contains the channel of

- instruments 61-80.

- 5	Instrument Channel String		ins05	80C	Setup. The string contains the channel of

- instruments 81-100.

- 6	Instrument Channel String		ins06	80C	Setup. The string contains the channel of

- instruments 101-120.

- 7	Instrument Channel String		ins07	80C	Setup. The string contains the channel of

- instruments 121-140.

- 8	Instrument Channel String		ins08	80C	Setup. The string contains the channel of

- instruments 141-160.

- 9	Instrument Channel String		ins09	80C	Setup. The string contains the channel of

- instruments 161-80.

- 10	Instrument Channel String		ins10	80C	Setup. The string contains the channel of

- instruments 181-200.

### Q.iMPACT Global Parameters Table

ControlNet Class Code: 83 hex

- There is one instance of the Global Parameters Table.

- PAC GLOBAL PARAMATERS SHARED DATA

Attr. 	Function/Value			Name	Format	Description

1	Max Purchased Channels		as101	US	Maximum Purchased Channels – up to 200

channels.

2	Max Purchased Material-Paths	as102	US	Maximum Purchased Channels – up to 1000

material-paths.

3	Database Master Terminal		as103	US	Setup. 1-20.  Q.iMPACT Terminal that

has the master version of the PAC

Database.

4	Update Indicator			as104	US	Calculated.

5	Purchased Upgrades Check  Digit	as105	US	Check Digit for Purchased Upgrades.

Updated whenever as101 or as102 change.

6	Spare				as106	20L	Spare.

7	Spare				as107	20L	Spare.

8	PAC XREF Status		as108	UC	PAC Database XREF Status.  0= Needs to

be built, 1= Built

9	PAC Global Data Base		as109	6C	 “GLOBL” indicates that Web pages update

Data Base globally.  “LOCAL” indicates that

Web pages update data bases locally.

10	PAC Weight Entry Method		as110	US	0 = Absolute Weight, 1=% of Capacity

11	Spare				as111	20L

12	Spare				as112	160C

## Q.iMPACT – Controller Communications HEAP Fields

A Controller communicates with the Q.iMPACT through ControlNet, ControlNet/Ethernet IP, or TCP/IP/Ethernet communications.

The Q.iMPACT uses the Shared Data fields in this section for building and processing communication messages to a Controller.  The Q.iMPACT stores these Shared Data fields in HEAP memory.  The Q.iMPACT only actively maintains these fields when it has a local Controller Interface.

The Controller sends commands to the Q.iMPACT using two Shared Data Messages.  The Controller sends the command with one message.   Then it requests the Q.iMPACT to send the response to the command in the second message.  The format of all command and response messages is the same for all flow meter and scale instruments.

The Controller addresses the command to an instrument.  It can only send one command to one instrument at a time.   The Q.iMPACT processes the commands immediately.  After sending a command to an instrument, the Controller must ask for the response to that command.  For most commands, the Q.iMPACT will already have the response to the command available when the Controller asks for it.  However, in the case that the response is not ready, the Q.iMPACT will reply with a “Command not Complete” status. This will most likely happen on commands to instruments located in remote Q.iMPACT terminals.  The controller must only ask for the response again after a short delay, and then it will receive the actual command status.

When the Controller needs to send commands to multiple instruments at once, the Controller must send the commands and request the responses for each instrument in strictly ordered pairs.  That is, it must send the command to instrument #1 and then ask for the response to instrument #1.  Then, it can send the command to instrument #2 and then get the response for instrument #2.

If the Controller does not maintain this precise request/response pairing, it may get the response matched to the wrong command.   In order to prevent this from happening, the Controller should always compare the first four bytes of the request and response messages.  They should always be identical for a matching request and response.  If they are not identical, there is a coding or communication error that must be corrected.

The “Master Reset” commands are exceptions.  The Q.iMPACT may send the command to multiple instruments with a single command.  It sends the “Master Reset – Cluster” command to all instruments in the cluster. The Master Reset commands put the Q.iMPACT instruments back in their initialized state.  The Controller must maintain the same request/response protocol for these commands.

### PAC Communications Command Data

ControlNet Class Code: 84 hex

- There is one instance of the PAC Command Data Object.

The Q.iMPACT stores these Shared Data fields in HEAP memory.  It only actively maintains these fields when it has a local Controller “Bridge” Interface.

A Controller sends a Shared Data Request Message to the Q.iMPACT to give a specific command to a specific instrument or Unit.  To send the message, the Controller Ladder program executes a MESSAGE instruction to “Set All Attributes” in the PAC Command “ACM00” Object in the Q.iMPACT.  The table below shows the PAC Command Object format.  The first four fields of this message are identical in this request message and in its matching response message.

- PAC Communications Command SHARED DATA

Attr. 	Function/Value			Name	Format	Description

- 1	Channel Number			acm01	UC	Instrument Channel for Command.

- 2	Message Sequence Number	acm02	UC	Message Sequence Number of Command.

- 3	Material Path Index		acm03	US	Material Path Index

4	Command Number		acm04	UC	“Command” Number.  Refer to “akn04”.

5	Overlap Feed Group		acm05	UC	Overlap Feed Group.  Refer to “akn05”.

6	Number of Overlapping Feeds	acm06	US	Number of Overlapping Feeds that are

simultaneously going into a Unit.  Refer to “akn06”.

7	Target Weight			acm07	F	Target Feed Weight.

8	Positive Tolerance		acm08	F	Positive Feed Tolerance.

9	Negative Tolerance		acm09	F	Negative Feed Tolerance.

10	Batch ID/ Display Message	acm10	40C	Batch ID from Host Controller that is used for

Data Collection Messages.   If there is a “^”

in the field, data following the “^” is used

as a Display Message for the Q.iMPACT

display.

### PAC Communications Command Status

ControlNet Class Code: 85 hex

- There is one instance of the PAC Command Status Object.

The Q.iMPACT stores these Shared Data fields in HEAP memory.  It only actively maintains these fields when it has a local Controller Interface.

The Q.iMPACT sends the PAC Command Status Shared Data to the Controller to give the immediate status for a command.  To get the Command Status Object, the Controller Ladder program must execute a MESSAGE instruction to “Get All Attributes” of the PAC Command Status “ACS00” Object in the Q.iMPACT.  The first four fields of the Command Status Object are identical to the Command Object in Request Message #2.  If these fields do not match in the two messages, the Controller will know that the response is not valid and must take corrective action.

- PAC Communications Command Status SHARED DATA

Attr. 	Function/Value			Name	Format	Description

- 1	Channel Number			acs01	UC	Instrument Channel for Command.

- Same as acm01.

- 2	Message Sequence Number	acs02	UC	Message Sequence Number of Command.

- Same as acm01.

- 3	Material Path Index		acs03	US	Material Path Index.  Same as acm01.

- 4	Command Number		acs04	UC	Command Number.  Same as acm01.

5	Command Status			acs05	UC	Command Status.  Refer to “arn05”.

6	Last Material Transfer Status	acs06	US	Last Material Transfer Status.  Refer to

“arn06”.

7	Transfer Status Qualifier Bits	acs07	L	Transfer Status Qualifier Bits.  Refer to

“arn07”.

8	Delivered Weight			acs08	F	Delivered Weight at Completion of Feed.

9	Deviation from Target Weight	acs09	F	Deviation from Target Weight

### Q.iMPACT->PLC Cyclic Communications Input Assembly

ControlNet Class Code: 86 hex

- There are 25 instances of the Q.iMPACT->PLC Input Assembly.  One instance is reserved for future growth.

The Q.iMPACT stores these Shared Data fields in HEAP memory.  It only actively maintains these fields when it has a local Controller Interface.

The Q.iMPACT sends the real-time process data to the Controller through the Q.iMPACT->Controller Cyclic Input Assembly message.  Each Q.iMPACT channel, representing either a weigh scale or flow meter, has a Q.iMPACT->Controller Input Assembly data block that is 20 bytes long.  The format of the block is the same for all channels.  The Q.iMPACT packs multiple instrument channels into one large Cyclic Input-to-Controller Assembly message.  24 instrument channels can fit snuggly into one 496-byte Class 1 message, with 16 bytes reserved for future growth.

This approach requires a separate Q.iMPACT Bridge terminal containing a ControlNet Interface Card for every 24 channels in the Q.iMPACT Cluster.  For example, to support a cluster with 96 channels, you would need four ControlNet Interface cards in the cluster in four separate Q.iMPACT terminals.  A Q.iMPACT cluster can contain up to 200 channels.

The Instrument Channel Configuration Object has the assignment of instrument channels to the Q.iMPACT Bridge terminal containing the ControlNet interface.  The following table shows the arrangement of instrument data within the larger assembly message for multiple instrument channels.

| Reserved |  | Data Byte 1 | Data Byte 2 | ………… | Data Byte 16 |
| --- | --- | --- | --- | --- | --- |
| Assembly Slot 1 |  | Data Byte 17 | Data Byte 18 | ………… | Data Byte 36 |
| Assembly Slot 2 |  | Data Byte 37 | Data Byte 38 |  | Data Byte 56 |
| Assembly Slot 23 |  | Data Byte 457 | Data Byte 458 |  | Data Byte 476 |
| Assembly Slot 24 |  | Data Byte 477 | Data Byte 478 | …………… | Data Byte 496 |

The Q.iMPACT normally updates data for the Cyclic Input-to-Controller Assembly once a second.   In case of critical events, such as, the completion of a material movement operation or an error, the Q.iMPACT updates the status immediately.  The Controller sets the rate at which it cyclically reads the Assembly.

The Cyclic Input-to-Controller Assembly data for each instrument channel is 20 bytes long.  Refer to the Process Table, which shows the format of the assembly.

- Q.iMPACT->Controller Cyclic Communications Input Assembly

Attr. 	Function/Value			Name	Format	Description

- 1	Reserved			aa101	16C	Reserved

- 2	Channel #1 PLC Input Assembly	aa102	20C	Channel #1 PLC Input Assembly

- 3	Channel #2 PLC Input Assembly	aa103	20C	Channel #2 PLC Input Assembly

- .	.	.			.	.	.

- .	.	.			.	.	.

- 25	Channel #24 PLC Input Assembly	aa125	20C	Channel #24 PLC Input Assembly

## PAC Processes

There are 12 PAC Process Identifiers within a Q.iMPACT, and, at most, only 12 PAC Processes can be active at a time within each Q.iMPACT.  There can be 240 PAC Process Identifiers in a Q.iMPACT cluster, but there are can be at most 200 active PAC Processes in a Q.iMPACT cluster.  When there are one to four scales and up to eight flow meters in the Q.iMPACT, the assignment of devices to processes within a Q.iMPACT is as follows:

Process 1 	Flow Meter K		|	Process 2	Flow Meter L
Process 3 	Flow Meter M		|	Process 4 	Flow Meter N

Process 5 	Flow Meter O		|	Process 6 	Flow Meter P

Process 7	Flow Meter Q		|	Process 8	Flow Meter R
Process 9 	Scale A 		|	Process 10	Scale B

Process 11	Scale C 		|	Process 12	Scale D

When there are no scales and up to 12 flow meters in the Q.iMPACT, the assignment of Processes to Flow Meters is as follows:

Process 1 	Flow Meter K		|	Process 2	Flow Meter L
Process 3 	Flow Meter M		|	Process 4 	Flow Meter N

Process 5 	Flow Meter O		|	Process 6 	Flow Meter P

Process 7	Flow Meter Q		|	Process 8	Flow Meter R
Process 9	Flow Meter S		|	Process 10	Flow Meter T
Process 11	Flow Meter U		|	Process 12	Flow Meter V

### PAC Process Command Table

A Controller sends a Shared Data Request Message to the Q.iMPACT to give a specific command to a specific instrument channel or Unit.  To send the message, the Controller Ladder program executes a MESSAGE instruction to “Set All Attributes” in the PAC Command “ACM00” Object in the Q.iMPACT with the ControlNet Interface.  A Command Dispatcher Task sends the host command to the Shared Data field for the specific PAC Process.

The Command Dispatcher directs the command to a specific Q.iMPACT and Process Identifier.   When PAC Process receives a command to begin moving material, the PAC Process uses the Material-Path Index field to find the appropriate Material-Path Table entry and the Unit Table entry.  The command includes the amount of material to be moved and the required tolerances. The PAC object retrieves the feed parameters from the Material-Path and Unit Tables.

The PAC Process initiates and controls the Material Transfer associated with the command.  The PAC Process operates once a second to update the PAC cutoff values based on the current the scale or flow meter status information.  The PAC Process interacts with the Scale and Flow Meter tasks to get the current status and to set cutoff values.

ControlNet Class Code: 87 hex

An instance of the PAC Commands Object may exist for each of 12 PAC Processes.  Each contains the numbered instance attributes.

- PAC PROCESS COMMAND SHARED DATA

- Where the ‘n’ is the instance representing the internal PAC Process number 1-9, A, B, C.

Attr. 	Function/Value			Name	Format	Description

- 1	Channel Number			akn01	UC	Channel Number for Command.

- 2	Message Sequence Number	akn02	UC	Message Sequence Number of Command.

- 3	Material Path Index		akn03	US	Material Path Index

4	Command Number		akn04	UC	“Command” Number

- None

- Start Material Transfer

- Start Material Transfer with Gross Weight Target.  Valid only for scale devices.

- Start Hand Add

- Acknowledge Material Transfer or Hand Add Complete

- Abort Material Transfer

- Reset Slow Step Timer

- Start Manual Mode

- Turn on FCE in Manual Mode

- Turn off FCE in Manual Mode

- Restart Auto Mode

- Complete Feed in Manual Mode

- Master Reset – Instrument Channel

- Report Last Status

- Master Reset – Cluster

- Validate Aggregate Secondary Feeds

5	Overlapping Feed Group		akn05	UC	Identifies which Primary and Secondary

Feed requests belong to a group of feeds

that make up an overlapping feed.  A value =

0 indicates that this is not an overlapping

feed.

6	Number of Overlapping Feeds	akn06	US	Number of Secondary Overlapping Feeds

that are simultaneously going into a Unit.  This field is only meaningful in a Primary Overlapped Feed command for a scale instrument.  The Material-Path must have a Gain-In-Weight feed.  The PAC turns on the FCE when it determines there will be enough time after the overlap completes to run the PAC algorithm with the scale.

7	Target Weight			akn07	F	Target Feed Weight.

8	Positive Tolerance		akn08	F	Positive Feed Tolerance.

9	Negative Tolerance		akn09	F	Negative Feed Tolerance.

10	Batch ID/ Display Message	akn10	40C	Batch ID from Host Controller that is used for

Data Collection Messages.   If there is a “~”

in the field, data following the “~” is used

as a Display Message for the Q.iMPACT

display.

### PAC Process Command Status Table

The Q.iMPACT sends PAC Command Status Shared Data to the Controller to give the immediate status for a command.  To get the Command Status Object, the Controller Ladder program must execute a MESSAGE instruction to “Get All Attributes” of the PAC Command Status “ACS00” Object in the Q.iMPACT.

A Command Dispatcher in the Q.iMPACT reads the status for the appropriate Q.iMPACT terminal and PAC process and returns that status to the host.  Each PAC Process maintains a its Command Status data in Shared Data.

ControlNet Class Code: 88 hex

An instance of the PAC Command Status Object may exist for each of 12 Local PAC Processes.  Each contains the numbered instance attributes.

- PAC PROCESS COMMAND STATUS SHARED DATA

Attr. 	Function/Value			Name	Format	Description

- 1	Channel Number			arn01	UC	Channel Number. Set on receiving a

- command.

- 2	Message Sequence Number	arn02	UC	Message Sequence Number. Set on

- receiving a command.

- 3	Material Path Index		arn03	US	Material Path Index of command.  Set on

- receiving a command.

- 4	Command Number		arn04	UC	Command Number of command.  Set on

- receiving an command.

5	Command Status			arn05	UC	“Command Status”.  Set immediately after

processing command.

- SUCCESS – Start Gain In Weight  Material Transfer Command Complete

- SUCCESS – Start Loss In Weight Material Transfer Command Complete

- SUCCESS – Start Flow Meter Material Transfer Complete

- SUCCESS – Start Manual Control Material Transfer Complete

- SUCCESS – Hand Add Command Complete

- SUCCESS –  Command Complete

- Command Not Complete – Request status again after a short delay

- ERROR – Communications Error

- ERROR – Invalid Channel Number

- ERROR – Invalid Command

- ERROR – Invalid Material-Path Table Index Number

- ERROR – Invalid Algorithm in Material-Path Table Entry

- ERROR – Invalid Feed Type in Material-Path Table Entry

- ERROR – Invalid Unit Table Index Number in Material-Path Table Entry

- ERROR – Invalid Gain In Weight Feed and Dump to Empty Algorithm Combination in Material-Path Table Entry

- ERROR – Source / Destination Mismatch in Material-Path Table Entry

- ERROR – Other Invalid Data in Material-Path Table Entry

- ERROR – Invalid Loss In Weight Feed in Material-Path Table Entry and Overlapping Feed Command

- ERROR – Invalid Data in Measuring Device Table Entry

- ERROR – Invalid Mode for Command.  e.g.  Controller is requesting start a new material transfer before last feed is complete or before the controller has acknowledged that the material transfer is complete.

- ERROR – Requested add amount too small

- ERROR – Requested add amount would bring unit over capacity

- ERROR – Unit Currently over Capacity

- ERROR – Unit Currently under Zero

- ERROR – Instrument Malfunction

- ERROR – Target Weight is Less Than Spill

- ERROR – Response Timeout

- ERROR – Too many overlapping feeds

- WARNING – Delayed start to feed due to overlapping feed

- WARNING – Abort ignored since time to complete was less than feed override time.

- ERROR - Invalid overlap group number.

- WARNING –Waiting For All Secondary Requests.

- WARNING – Waiting For Measruing Device Stability.

- ERROR – Not Enough Material.

- ERROR – Device Not  Calibrated Properly.

6	Last Material Transfer Status	arn06	US	Last Material Transfer Status.  Set at

completion of feed.

- Successful Material Transfer – K1, K2 parameters updated

- Successful Material Transfer – Spill Only

- Successful Material Transfer – Dump to Empty

- Successful Hand Add

- Material Transfer Complete – K1, K2 parameters NOT updated

- Material Transfer Complete – Parameters Reset

- Material Transfer Complete with Manual Operation

- Failed – Unstable Measuring Device

- Failed – Overlapping Feed Error Corrupted Flow

- Failed – Erratic Flow Error

- Failed – Low Flow Error

- Failed – High Flow Rate Alarm Error

- Failed – Communication Error

- Failed – Instrument Error

- Failed – Unit Capacity Error

- Failed – Predictive Algorithm Error

- Failed – Material Transfer with Manual Operation

- Failed – Amount of material transferred did not match in source and destination

- Failed – Controller Aborted Material Transfer

- Failed – Controller Reset Channel

- Failed – Controller Reset Unit

- Failed – Controller Reset Cluster

- Failed – Slow Step Timer Timeout

- Failed – Secondary Requests Timeout

- Failed – Power Failure During Feed

- Failed – Command Failed, Transfer Not Started

- Material Transfer In Progress

7	Transfer Status Qualifier Bits	arn07	L	Transfer Status Qualifier Bits. Set at

completion of feed.

- Over Tolerance

- Under Tolerance

- Power Failure During Feed

3-15 Reserved

8	Delivered Weight			arn08	F	Delivered Weight.  Set at completion of

Feed.

9	Deviation from Target Weight	arn09	F	Deviation from Target Weight.  Set at

completion of feed.

### PAC Process Table

The Q.iMPACT runs the PAC Processes once a second.  When it runs, the PAC Process sets its current status in the Process Table Shared Data, as necessary.  In case of critical events, such as, the completion of a material movement operation or an error, the Q.iMPACT updates the status immediately.   The PAC Process also reports the first 8 fields in this table in the Q.iMPACT->Controller Cyclic Input Assembly.

ControlNet Class Code: 89 hex

An instance of the PAC Process Table Object may exist for each of 12 PAC Processes.  Each contains the numbered instance attributes.

- PAC PROCESS TABLE SHARED DATA

- Where the ‘n’ is the instance representing the internal PAC Process number 1-9, A, B, C.

Attr. 	Function/Value			Name	Format	Description

The Q.iMPACT sends the first 8 fields to Controller via cyclic communications.

- 1	Channel Number			ptn01	UC	Channel Number (1-200)

2	Real-time Status 1		ptn02	UC	Status 1 bit assignment.

- Communications Data Integrity Bit alternates Polarity Each Second

- Instrument Data Integrity OK

- Scale Over Capacity

- Scale Under Zero

- Scale Motion

- Material Transfer Cycle Active

- Discrete IO Final Control Element                0 = Off , 1 = On

- Waiting for Controller to Acknowledge Last Material Transfer/Hand Add complete

3	Real-time Status 2		ptn03	US	Status 2 bit assignment.

- Feed Type Bits1&2.

- Feed Type Bit  1&2.  0 = Gain In Weight, 1=Loss In Weight, 2 = Flow Meter, 3 = Hand Add.

- Manual (Not Auto) Mode

- Gross Weight Feed

- Feed Override Active – external logic inhibited from removing permissive

- Feed Failed

- Communication Error

- Device Stability Warning

- Very Unstable Device

- Too High or Too Low Flow at Cutoff

- Three Times Average Flow at Cutoff

- Fast Feed Rate Alarm

- Wait for All Overlap Requests

- Waiting to Start Primary Overlapped Feed

- Primary Overlapped Feed In Progress

- Secondary Overlapped Feed.

4	Feed Weight			ptn04	F	Feed Weight.  During feed, this field is the

Net Accumulated Weight.  At completion of

feed, this field is the Delivered Weight.  This

field is reset to zero at the beginning of a

feed.

5	Gross Weight			ptn05	F	Gross Weight.  For scales, this field the

gross weight.  For flow meters, this field is

the same as Feed Weight.

6	Rate				ptn06	F	Rate of change of weight in primary units per

second.

7	Slow Step Timer			ptn07	US	“Time until Slow Step Timer Expires” in

Seconds = Slow Step Factor * (target /

average flow)

8	Estimated Time to Complete	ptn08	US	“Estimated Time to Complete” in Seconds

The Q.iMPACT does NOT send the following fields to Controller via cyclic communications.

- 9	Current Material Path Index	ptn09	US	Current Material Path Index.

- 10	Current Command Number	ptn10	UC	Current Command Being Processed.

- 11	Current Overlap Group		ptn11	UC	Current Overlap Feed Group.

- 12	Current # Overlapping Feeds	ptn12	US	Current # Overlapping Feeds.

13	Current Target Weight		ptn13	F	Current Target Feed Weight.

14	Current Positive Tolerance		ptn14	F	Current Positive Feed Tolerance.

15	Current Negative Tolerance	ptn15	F	Current Negative Feed Tolerance.

16	Batch ID/ Display Message	ptn16	40C	Batch ID from Host Controller that is used for

Data Collection Messages.   If there is a “^”

in the field, data following the “^” is used

as a Display Message for the Q.iMPACT

display.

17	Feed Time			ptn17	Long	Feed time in seconds. PAC sets this current

value once a second during feed and sets

total time at completion of feed.

18	Weight Units			ptn18	3C	ASCII descriptor for weight units. “lb”, “kg”,

”g”,  or “t”

19	Diagnostic Status			ptn19	Long	PAC sets feed diagnostic status word during

feed cycle.  At completion of feed, the limits

values reflect status of feed.  Refer to qnn31

for bit meaning assignments.

20	Feed Start Time			ptn20	L	Time in seconds from 1970 GMT

21	Unit Weight at Start		ptn21	D	Unit Weight at Start of Feed

22	Unit Weight at Completion		ptn22	D	Unit Weight at Completion of Feed

### PAC Process History Data

PAC makes an entry in the Process History record at the completion of each Material Transfer for the PAC Process.   There are two data records for each PAC Process. These Shared Data fields allow the PAC Historian to retrieve the history record for each process.

ControlNet Class Code: 91 hex

PROCESS HISTORY SHARED DATA	(External Read Only)

Attr. 	Function/Value			Name	Format	Description

1	History Record 1			hst01	400C	Process History Record #1, Process #1.

Contains the following fields in ASCII format

Separated by ‘^’ characters in the following

order:

sequenceNumber, channelNumber,

materialPath, command, status,

lastMaterialTransferStatus,statusQualifier,

deliveredWeight,  spill_deviationFromTarget ,

realtimeStatus1, realtimeStatus2,

slowStepTimer, estimatedTimeComplete,

overlapGroup,

numSecondaryOverlappingFeeds,

targetWeight, positiveTolerance,

negativeTolerance, displayMessage,

feedTime, weightUnits,  startTime,

startWeight, endWeight .

2	History Record 2			hst02	400C	Process History Record #2, Process #1.

Contains the measuring device table

entry, material path table entry table fields in

ASCII format separated by ‘^’ characters in

sequential order.

3	History Record 3			hst03	400C	Process History Record #1, Process #2.

4	History Record 4			hst04	400C	Process History Record #2, Process #2.

.           .				.	.	.

.           .				.	.	.

23	History Record 23	 	hst23	400C	Process History Record #1, Process #12.

24	History Record 25		hst24	400C	Process History Record #2, Process #12.

25	History Record Sequence #	hst25	12L	Sequence # for each Process.

## Measuring Device Channel  Table

Each Measuring Device Channel Table entry contains several static parameters associated with flow control for each measuring device.  A Measuring Device is a Scale or Flow Meter.  The System Engineer initially sets up these parameters at system installation and only changes them to “fine tune” or to reconfigure the system.   Some parameters are only applicable to scale devices.

ControlNet Class Code: 8A hex

An instance of the Measuring Device Channel Table Object may exist for each of 200 instruments in a Q.iMPACT cluster.  Each contains the numbered instance attributes.

- MEASURING DEVICE CHANNEL TABLE SHARED DATA

- Where the Hex number  ‘nn’ is the instance representing the Measuring Device 1-200.

Attr. 	Function/Value			Name	Format	Description

1	Measuring Device Assigned	unn01	UC	Setup. 1=Yes, 0=No

2	Q.iMPACT Terminal		unn02	C	Setup.  0, 1-9, A-K.  Cluster Terminal

containing Measuring Device.

3	Measuring Device		unn03	C	Setup.  Scale A – D or Flow Meter K – V.

4	Stable Scale Time		unn04	US	Setup. Number of seconds to wait for a

stable scale  reading before returning an

“Unstable Scale” failure status.  Range is 5-

10 seconds.

5	Reserved			unn05	US	Reserved.

- 6		Feed Override Time		unn06	US	Setup. Time in seconds that any external

- logic (such as Slow Stem Timer or an

- operator changing modes) will be inhibited

- from removing the permissive on the

- enabling logic that is controlling the Final

- Control Element (FCE). This is used to

- prevent  something other than the Fast-Cut-

- Off from closing the FCE, which might cause

- the PAC to develop erroneous data for

- subsequent update of its constants.

- The PAC sets a status indicating that the

- Material Transfer is within the feed override

- Time. Default = 20 seconds

- 7		Minimum Slow Step Time		unn07	US	Setup.  The PAC algorithm computes a “step

- slow” value.  If the computed value is less

- than the minimum specified here,  PAC uses

- the minimum instead. Typical range is 30-60

- seconds.

- 9		Overlap Feed Alone Time		unn09	US	Setup. Time that an overlapped feed must

- fed alone before cutoff, in seconds. Typically

- 10-20 seconds.

- 10		Overlap Time Tolerance		unn10	US	Setup. Additional time in seconds allowed for

- an overlapping feed to complete.

- 11		Zero Flow Threshold		unn11	Float	Setup. Flow rate below which the system

- assumes zero flow. Typically sized 0.001%-

- 0.01% of Max Scale Capacity/second.  In a

- “dump to empty”,  the PAC uses this value to

- determine when a “dump to empty” is

- complete.  After this point, PAC keeps the

- FCE energized until the “drain time“ expires.

- When this value < 0, PAC task uses absolute

- Value for drain timer for and aborts drain

- Timer at zero flow.

- 12		Unstable Device Flow Threshold	unn12	Float	Setup. Flow rate above which PAC

- generates a “Noisy Scale” condition when

- waiting for a stable scale reading. Typically

- sized 0.005%-0.05% of Max Scale Capacity

- per second. Must be larger than Minimum

- Flow.

- 13		Minimum Add			unn13	Float	Setup.  The smallest amount of material the

- system will attempt to add. Typically 0.01%-

- 1% of Max Scale Capacity.

- 14		Maximum Unit Size		unn14	Float	Setup. Over capacity point of a vessel, in

- weight units.

- 15		Dump Trip Point			unn15	Float	Setup. Level below which PAC considers

- Dump To Empty complete. Typically,

- 0.001%-0.01% of Max Unit Size. Must be

- less than smallest batch size for the Unit.

- 16		Min PAC Flow Rate Threshold	unn16	Float	Setup.  PAC starts to apply the predictive

- algorithm when measured flow exceeds this

- value.  Typically, 0.1% of Max Unit Size.

- Must be larger than Minimum Flow.  PAC

- sets the setpoint target = SP – Spill until the

- the flow rate reaches the this value.

- 19		Reserved			unn19	Float	Reserved.

- 30		Current Zero			unn30	Float	Calculated.  After a “Dump to Empty”

- operation, PAC sets this value to actual

- scale weight when it algorithmically detects

- the zero.  It helps track a “heel” buildup in

- the vessel.

## Measuring Device Channel Defaults Table

The Measuring Device Channel Defaults Table entry contains default parameters associated with the Measuring Device Channel Table entries.  The Measuring Device Channel Table Setup Web Pages use these defaults to dynamically change the default values for the Measuring Device Channel Table.   The Q.iMPACT stores them in BRAM.

ControlNet Class Code: 8B hex

- MEASURING DEVICE CHANNEL DEFAULT TABLE SHARED DATA

Attr. 	Function/Value			Name	Format	Description

1	Default Unit Assigned		dfu01	UC 	Same As Measuring Device Channel Table.

2	Default Q.iMPACT Terminal	dfu02	C 	Same As Measuring Device Channel Table.

3	Default Weigh Scale		dfu03	C 	Same As Measuring Device Channel Table.

4	Default Stable Scale Time 		dfu04	US 	Same As Measuring Device Channel Table.

5	Default Reserved			dfu05	US 	Same As Measuring Device Channel Table.

6	Default Feed Override Time	dfu06	US 	Same As Measuring Device Channel Table.

7	Default Minimum Slow Step Time	dfu07	US 	Same As Measuring Device Channel Table.

9	Default Overlap Feed Alone Time	dfu09	US 	Same As Measuring Device Channel Table.

10	Default Overlap Time Tolerance	dfu10	US 	Same As Measuring Device Channel Table.

11	Default Minimum Flow		dfu11	Float 	Same As Measuring Device Channel Table.

12	Default Maximum Flow		dfu12	Float 	Same As Measuring Device Channel Table.

13	Default Minimum Add		dfu13	Float 	Same As Measuring Device Channel Table.

14	Default Maximum Unit Size	dfu14	Float 	Same As Measuring Device Channel Table.

15	Default Dump Trip Point		dfu15	Float 	Same As Measuring Device Channel Table.

16	Default Min PAC Flow Rate	dfu16	Float 	Same As Measuring Device Channel Table.

19	Default Reserved			dfu19	Float 	Same As Measuring Device Channel Table.

30	Last Zero			dfu30	Float 	Same As Measuring Device Channel Table.

## Material-Path (i.e. Raw Material) Table

The Material-Path Table defines the routing and flow control of each Raw Material through each specific Measuring Device and Unit in the Q.iMPACT.  Each Q.iMPACT may contain the definition for all the Material-Paths for the entire Q.iMPACT cluster, but it must only necessarily have the data associated with local Material-Paths.  There may be up to 1000 Material-Paths in the Q.iMPACT cluster.

The Material-Path entry contains both static Setup parameters and dynamic PAC parameters. The System Engineer sets up the static parameters that are configuration and static processing parameters for the Material-Paths.  This data only changes when the System Engineer is “fine-tuning” the system or is reconfiguring the factory.

The Dynamic PAC parameters are variables that dynamically control the feed cutoffs.  These variables are set initially to a default value, but their values may change after every feed cycle.  Examples of these parameters are the PAC algorithm cutoff “constants.”  The PAC algorithm automatically calculates the optimum values for these variables after a minimum number of feeds. Other dynamic parameters include the counts for good and bad feeds.  Only the local Q.iMPACT has an up-to-date copy of the dynamic parameters.

ControlNet Class Code: 8C hex

An instance of the Material-Path Object may exist for each of 1000 Material-Paths.  Each contains the numbered instance attributes.

MATERIAL-PATH TABLE SHARED DATA	(External Read/Write)

Where the ‘nn’ is the instance representing the internal Unit number 1-1000, in base 36 number representation.

Attr. 	Function/Value			Name	Format	Description

1	Material-Path Assigned		qnn01	UC	Setup. 1=Yes, 0=No

2	Instrument Channel		qnn02	UC	Setup. Channel 1 to 200 is the Q.iMPACT

logical measuring device channel in the

cluster that processes this Material-Path.

3	Text Descriptor			qnn03	19C	Text Descriptor for Material-Path ID

5	PAC Algorithm			qnn05	UC	Setup.  PAC Algorithm.

0 = Spill Only,

1 = K1,

2 = K2,

3 = Dump To Empty.

6	Feeder Type			qnn06	UC	Setup.   Feeder Type

- Gain In Weight,

- Loss In Weight,

- Flow Meter

- Hand Add

7	Slow Step Timer Factor		qnn07	F	Setup.  The Slow Step Timer is defined

Factor*( target / average flow ).  The factor

is normally set as 1.5, but is adjustable on a

material basis.

If SST Factor< 0, used absolute value for

SST Factor, but generate alarm only

When SST expires.

9	Minimum Open Time		qnn09	US	Setup.  Time in seconds that spill

compensation is not applied immediately

following control device opening. If = 0, no

minimum.   It is the minimum time in

seconds that a feed must be active

before it is considered “successful” and

will have its PAC parameters updated.

This check guarantees that the flow rate

is valid before PAC updates the PAC

parameters for a Material Path.

10	Drain Time			qnn10	US	Setup.  Time in seconds that the system will

wait for material to drain into a vessel after

feed is complete and before testing for

material delivery tolerance.

11	Average Flow Rate Low Limit	qnn11	Float	Setup. Lower alarm limit for the Average

Flow Rate “A”. Typically set at 50% of the

flow rate, in weight or volume units/second.

12	Average Flow Rate High Limit	qnn12	Float	Setup. Upper alarm limit for the Average

Flow Rate “A”. Typically set at 150% of the

flow rate, in weight or volume units/second.

13	Average Spill Low Limit		qnn13	Float	Setup.  Lower alarm limit for the Average

Spill “AA”. Typically set at 50% of the

average spill, in weight or volume units.

14	Average Spill High Limit		qnn14	Float	Setup.  Upper alarm limit for the Average

Spill “AA”. Typically set at 150% of the

average spill, in weight or volume units.

15	Algorithm Update Parameter	qnn15	Float	Setup. PAC uses this value in calculation of

Average Flow Rate “A”, Average Spill “AA”,

and Cutoff Constants “B”, “BB”, and “C”.

It controls how quickly the system

responds to a change in operating

conditions. Range is 0.0–1.0. Smaller values

(0.1–0.3) should be used for systems that

have few changes, while larger values (0.6-

0.8) are used for systems that change

frequently. Default value is 0.2.

16	Flow Rate Filter Sample Period	qnn16	UC	Setup.  This value determines specifies the

period of time, from 1 to 60 seconds, over

which the rate is calculated.   For lower

values,  PAC responds more quickly to

changes in rate.  For larger values, the rate

value changes more smoothly.

17	Destination Channel		qnn17	UC	Setup. Destination channel.  Range 0 –

200. If the Destination Unit = 0, PAC

disables Unit Verification and Batch Count

Checking. At beginning of the feed, the PAC

finds the weight in the destination channel to

determine if this Material Feed will cause an overflow in the destination channel.  At the end of the feedings, the PAC determines if the loss in weight in the source channel matches the gain in weight in the destination unit.  Note that there may be multiple concurrent feeds into the destination channel. Destination Channel = 255 indicates that the Destination is outside of this cluster.

18	Max Flow Rate Alarm Value.	qnn18	Float	Setup.  Flow rates above this value generate

and alarm and terminate the feed.  Value = 0

turns off alarm checking.

19	Average Flow Rate “A”		qnn19	Float	Calculated/ Initialized in Setup. Average

Flow Rate “A”, in weight or volume units /

second.

20	Average Spill “AA”		qnn20	Float	Calculated/ Initialized in Setup. Average Spill

“AA”, in weight or 	volume units.

21	Cutoff Constant K1		qnn21	Float	Calculated.  First-order cutoff parameter

used in spill calculation.

22	Cutoff Constant K2		qnn22	Float	Calculated. Second-order cutoff parameter

used in spill calculation.

23	Flow at Last Cutoff “Q”		qnn23	Float	Calculated. Material flow rate at the cutoff of

the last feed cycle, in weight or volume units

/ second.

24	Spill at Last Cutoff		qnn24	Float	Calculated. Material spill following the last

feed cycle, in weight or volume units.

25	Cutoff Constant B			qnn25	Float	Calculated.  PAC uses “B” in K2 model. It is

the “Average Flow Squared”.

26	Cutoff Constant BB		qnn26	Float	Calculated.  PAC uses “BB” in K2 model.  It

is the “Average Flow * Spill”.

27	Cutoff Constant C			qnn27	Float	Calculated.  PAC uses “C” in K2 model.  It is

the “Average Flow Cubed”.

28	Consecutive Good Feed Counter	qnn28	Long	Calculated. Counts consecutive good feeds.

29	Total Good Feed Counter		qnn29	Long	Calculated. Counts total good feeds.

30	Total Bad Feed Counter		qnn30	US	Calculated. Counts the total bad feeds.

Feeds aborted early in the Material Transfer

cycle are not counted as bad feeds.

31	Diagnostic Status Bits		qnn31	Long	Calculated. PAC sets feed diagnostic status

word at  completion of last feed.  A bit SET

means:

- OK to process requested feed

- Minimum Open Time passed OK

- Minimum Feed Time passed  OK

- Predictive algorithm started  OK

- Predictive algorithm completed OK

- Feed Complete & Waiting for Stable Scale

- Over Tolerance

- Under Tolerance

- Flow Rate within limits

- Flow rate within 50% of range

- Flow rate within 75% of range

- Spill within limits

- Spill within 50% of range

- Spill within 75% of range

- Primary Feed Alone

- Awaiting Start Stability

- Dump Trip Point Passed

- K1 K2 Limit Violation

(Validating Aggregate Feed)

- Checking Fast Flow Rate Alarm

- Hand Add Complete Detected

- Feed Aborted Awaiting Drain

- Power Failure During Feed

- Configuration Reset During Feed

- Look_Back_Time too small

( Entered Setup During Feed )

- ETC Error on Last Primary Feed

- – 31 Reserved

32	Last Material Transfer Status	qnn32	UC	Calculated. Last Material Transfer Status is

set at completion of feed.  Refer to arn06.

33	Cutoff Constant A_hat		qnn33	Float	Calculated.  PAC uses in K2 model.

34	Cutoff Constant AA_hat		qnn34	Float	Calculated.  PAC uses in K2 model.

## Material-Path Defaults Table

The Material-Path Table defines the default values used in setting up the Material Path Table through the Web Pages.   The operator can dynamically change the defaults.  The Q.iMPACT stores them in BRAM.

ControlNet Class Code: 8D hex

MATERIAL-PATH DEFAULTS SHARED DATA	(External Read/Write)

Attr. 	Function/Value			Name	Format	Description

1	Default Material-Path Assigned	dfq01	UC 	Same as Material Path Table.

2	Default Instrument Channel	dfq02	UC 	Same as Material Path Table.

3	Default Text Descriptor		dfq03	19C 	Same as Material Path Table.

5	Default PAC Algorithm		dfq05	UC 	Same as Material Path Table.

6	Default Feeder Type		dfq06	UC 	Same as Material Path Table.

7	Default Slow Step Timer Factor	dfq07	F 	Same as Material Path Table.

9	Default Minimum Open Time	dfq09	US 	Same as Material Path Table.

10	Default Drain Time		dfq10	US 	Same as Material Path Table.

11	Default Flow Rate Low Limit	dfq11	Float 	Same as Material Path Table.

12	Default Flow Rate High Limit	dfq12	Float 	Same as Material Path Table.

13	Default Average Spill Low Limit	dfq13	Float 	Same as Material Path Table.

14	Default Average Spill High Limit	dfq14	Float 	Same as Material Path Table.

15	Default Update Filter Constant	dfq15	Float 	Same as Material Path Table.

16	Default Flow Rate Filter Constant	dfq16	UC 	Same as Material Path Table.

17	Default Destination Unit		dfq17	UC 	Same as Material Path Table.

18	Max Flow Rate Alarm Value.	dfq18	Float 	Same as Material Path Table.

19	Default Average Flow Rate “A”	dfq19	Float 	Same as Material Path Table.

20	Default Average Spill “AA”		dfq20	Float 	Same as Material Path Table.

21	Default Cutoff Constant K1		dfq21	Float 	Same as Material Path Table.

22	Default Cutoff Constant K2		dfq22	Float 	Same as Material Path Table.

23	Default Flow at Last Cutoff “Q”	dfq23	Float 	Same as Material Path Table.

24	Default Spill at Last Cutoff		dfq24	Float 	Same as Material Path Table.

25	Default Cutoff Constant B		dfq25	Float 	Same as Material Path Table.

26	Default Cutoff Constant BB	dfq26	Float 	Same as Material Path Table.

27	Default Cutoff Constant C		dfq27	Float 	Same as Material Path Table.
28	Default Good Feed Counter	dfq28	Long 	Same as Material Path Table.

29	Default Total Good Feed Counter	dfq29	Long 	Same as Material Path Table.

30	Default Total Bad Feed Counter	dfq30	US 	Same as Material Path Table.

31	Default Diagnostic Status Bits	dfq31	Long 	Same as Material Path Table.

32	Default Material Transfer Status	dfq32	UC 	Same as Material Path Table.

33	Cutoff Constant A_hat		dfq33	Float	Same as Material Path Table.

34	Cutoff Constant AA_hat		dfq34	Float	Same as Material Path Table.

## Flow Meter

### Flow Meter Heap Weight Object

ControlNet Class Code: 8E hex

An instance of the Flow Meter Weight Object may exist for each of 12 Flow Meters.  Each contains the numbered instance attributes.

FLOW METER WEIGHT SHARED DATA	(External Read Only)

Where the ‘n’ is the instance representing the internal flow meter number 1-C.

Attr. 	Function/Value			Name	Format	Description

1	Cumulative Weight		fwn01	13C	Cumulative Weight since start of feed,

rounded to nearest increment, in

ASCII string format.

- 3	Displayed Weight Units		fwn03	3C	3 alphanumeric (lb pounds, kg kilograms, grams, t metric tons)

7	Displayed Rate Period		fwn07	C	1 alphanumeric  (No, Sec, Min, Hour)

8	Displayed Rate			fwn08	13C	Rate of change of weight in string format.

9	Displayed Diagnostic Counts	fwn09	13C	Raw flow meter counts in string format.

10	Rounded Cumulative Weight	fwn10	F	Cumulative weight in float format.

rounded to nearest increment size.

14	Rate				fwn14	F	Rate of change of weight in single float

format.

15	Flow Meter State			fwn15	UC	0 = disabled, 1 = normal weight processing,

2 = diagnostic, 3 = calibration, 5 = error.

16	Masked Status Register		fwn16	US	Value of the Status Register on the Flow

Meter PCB.  Masked to show changes to

bit 0 (limit reached), bit 1(max rate

exceeded), and bit 4 (output state) only.

17	Fine Cumulative Weight		fwn17	F	Cumulative weight in float format.

18	Accumulative Flow Register	fwn18	L	Value of the Accumulative Flow Register on

Flow Meter PCB.

19	Unmasked Status Register		fwn19	US	Value of the Status Register on the Flow

Meter PCB.

### Flow Meter Calibration Object

ControlNet Class Code: 8F hex

An instance of the Flow Meter Calibration Object may exist for each of 12 flow meters.  Each contains the numbered instance attributes.

- FLOW METER CALIBRATION PARAMETERS STORED IN STATIC RAM

- Where the ‘n’ is the instance representing the internal flow meter number 1-C.

Attr. 	Function/Value			Name	Format	Description

1	Flow Meter ID			fsn01	20C	Flow Meter Identifier Text String.

2	Increment Size			fsn02	D	Increment size for rounding weight

3	Units				fsn03	UC	1=pounds, 2=kilograms, 3=grams,

4=metric tons

4	Command			fsn04	UC	Command

- None.  Flow Meter driver sets command field to “0” after completing command.

- Start Feed

- Abort Feed

- Set New Target for current feed.

- Load New Setup in Flow Meter Card

- Turn-on Discrete Output

- Turn-off Discrete Output

- Continue feed counting with manual control of discrete output

- Resume Feed (after abort)

- Start feed counting from 0 with manual control of discrete output.

5	Target Weight			fsn05	F	Target Amount to Feed

6	Maximum Rate			fsn06	F	Maximum Flow Rate

7	Rate Time Units			fsn07	C	No, Sec, Min, Hour

- 8	Flow Meter ‘K’ Factor		fsn08	US	Flow Meter ‘K’ Factor in pulses per unit (see fsn09 for unit selection)

9	Flow Meter ‘K’ Factor Units	fsn09	UC	0 = pulses/liter, 1 = pulses/cc,

2 = pulses/gallon, 3 = pulses/pound,

4 = pulses/kg, 5 = pulses/g

10	Accum. Flow Conversion Factor	fsn10	F	Multiplier used to convert the accumulated

flow register value to cumulative weight.

11	Flow Meter ‘M’ Factor		fsn11	US	‘M’ Factor for scaling the accumulated

flow register.

12	Flow Meter ‘R’ Factor		fsn12	US	‘R’ Factor for scaling the rate register.

13	Rate Filtering Window		fsn13	US	Sliding Window that is the number of

Seconds (0-60) over which the rate

Is averaged.

14	Capacity				fsn14	F	Capacity of Flow Meter counter.

# CHANGE HISTORY

    - Added Jag32 – Last Print Message.

Changed Template Length to 480.

Reduced # of Connections (cnn) to 18.

Increased # literals (lit) to 30.

Added s_24a – PLC Online Flag.

    - Changed length of net01 to 13c,  Ethernet address is stored as an ASCII string.

    - Comments only for csn21 and abn02.

    - Changed length of lad02 to 400.

    - Added jag33 through jag36.  JagX description fields

Added jag37.  Indirect access counter.

Added jag40 through jag49.  Indirect access pointers

01-13-2000 Added bd040 through bd049.  Alias names for jag40 through jag49.

02-17-2000 Added net30, net31, and net32.  Web Browser pages.

02-24-2000 Renamed bd006-bd007 to analog board part numbers.

Renamed bd008-bd009 to DigiTOL part number.

Renamed bd010-bd011 to idnet board part number.

03-20-2000 Added descriptions for t_60b and t_60c. Changed description for /emn02.

03-20-2000 Changed net28 to an unsigned short.

Added several spares: cs155-cs161;jag50-jag59;lit30-lit50;emn04-emn05

03-29-2000 Added lg700-lg900 for future use.

03-31-2000 Added restart network  t_606

    - Changed descriptions for net11, net12, net21, net22. Added s_25a  through s_25f.  Changed wsn14, and wsn20 through wsn24 from long to float.

    - Added jag38 as power cycle counter. Changed spares jag54-jag56 to unsigned short. Changed

Spares jag57-jag59 to doubles.	Changed spare csn59 to double.  Changed spare csn60 to US.

05-25-2000  Changed description of t_6da from Basic Execution Stopped to Basic Execution Running.

Added t_6df – Prompt List Running.

    - Added bd030 through bd036.  Deleted /_safl, /_sbfl, /_scfl, /_sdfl, /_sefl.

Corrected definition of csn85 to csn91 to UC.

Corrected definition of trn85 to trn96 to UC.

Corrected definition of spn85 to spn88 to UC.

Corrected definition of jag86 to jag91 to UC.

Corrected definition of srn85 to UC.

Corrected definition of abc85 to abc89 to UC.

Corrected definition of abn85 to UC.

Corrected definition of bas85 to bas89 to UC.

Corrected definition of rm101 to rm109 to UC.

Added bd014 and jag50.

Changed length of bd013 to 24.

06-07-00     Added bd050, bd051, bd052, bd053, jag50, jag51, jag52, jag53, jag60, jag61, jag62, jag63,

jag64, jag65, jag66, jag67, jag92.

Corrected weight fields in wtn00 and wsn00 from 12C to 13C.

Corrected var01 … from 45C to 47C

Corrected bd005 to bd012 from 12C to 13C.

Corrected csn02 from 6C to 7C.

Corrected csn18 from 8C to 9C.

Corrected csn27 from 11C to 12C.

Corrected csn30 from 11C to 12C.

Corrected spn01 from 8C to 9C.

Corrected jag01 from 2C to 3C.

Corrected jag07 from 8C to 9C.

Corrected jag08 from 8C to 9C.

Corrected jag11 from 12C to 13C.

Corrected jag12 from 12C to 13C.

Corrected jag18 from 10C to 11C.

Corrected jag19 from 11C to 12C.

Corrected jag20 from 11C to 12C.

Corrected jag21 from 10C to 11C.

Corrected scu11 from 8C to 9C.

Corrected scu12 from 8C to 9C.

Corrected sr109 from 6C to 7C.

Changed cnn18 to cnn12.

Corrected ce128 from 12 to 13C.

  - Added ControlNet Class Codes and Attribute assignments.

  - Made jag64 the post request for Alibi Memory searches.

  - Made editing changes for readability.

  - Added t_65c through t_65f

08-01-0	   JagXtreme Revision 0 Release

  - Added PrintingInProgress bits – s_264, s_26c, s_27c, s_28c, s_2fc; Added definitions for

Csn55 and csn56; Added t_60d

  - Added t_60e and t_60f; Corrected definition of wsn25

  - Added bd037, Changed CapturePowerCellCounts title to Capture Raw Counts.

  - Changed description of jag65 to battery replacement text string.

  - Changed AlarmOut from t_60f to s_2be.

09-05-0     Added bd015.

  - Added bx101-bx198 and bx201-bx298.

  - Added Autotune Variable description to bx101 – bx104.   Added definition for Jag57.

  - Corrected definition of jag54 – jag56 to double.

Corrected definition of jag57 - jag59 to unsigned short.

  - Added bd038 and s_60f.

  - Added t_61e and t_61f.

- 01-05-01   Added lg801, lg802, lg901, lg902.  Modified abc85.

    - Added PAC Shared Data

    - bd022, bd023,bd054,wt122, wt123,wt124,jag58, jag59, jag66. Changed definition of ce127. Made

extensive modification to PAC data definition.

    - Added bd055, jag67.   Removed qnn08.  Added Hand Add to qnn06.

    - Change qnn02 to instrument channel.

Delete qnn04.

Delete qnn18.

Change Process Number assignment for Scales and flow meters.

Change ins01 to contain device number instead of process number.

Increase csn18 and fsn01 from 9 to 20 characters

Change jag66 to 3 characters and jag67 to 4 characters.

    - Added Unit Table Defaults Shared Data.

Added Material-Path Defaults Shared Data.

Added qnn33. Changed to 19C.

    - Modified commands and statuses in acm04, acs05, akn04, and arn05.

    - Added fwn18 and fwn19.  Modified fsn10 usage and number type to float.  fsn04 to fsn06 are now spares.

    - Modified definitions for jag54, jag55, and qnn16.  Removed bx201-bx298 and nde01-nde10 for

PAC only; Modified definitions for fwn00 and fsn00.

04-16-01   Added bd060, bd061, and bd062.

    - Changed fwn18 from and unsigned short to a long.  Changed fwn16 from unsigned char to

unsigned short and changed its function.

    - Changed unn03 and dfu03 from ‘UC’ to ‘c’.

    - Updated description field for jag66.

    - Updated description of qnn06.  Modified description of acm10, akn10, & ptn16 to add Batch ID.  Added ptn20, ptn21, ptn22.

    - Added “rate alarm” status bits to aan03 and ptn03.

    - Added unn19 and dfu19.  Removed Manual Control from feed type.

    - Added new Command Statuses to arn05, added new Last Material Statuses to arn06.

    - Added status code 13 to ptn03.

    - Added command code 4 to fsn04 and added new fsn06.

    - Assigned jag56 as Mode Change Request and changed type from D to UC.

    - Added fsn13.  Added/clarified status bits in aan03.  Clarified definitions in akno5 & akn06.

Added status values to arn05 & arn06.

    - Added command 8 to fsn04.

06-20-01 Changed Unit Table to Measuring Device Channel Table, and modified table.  Deleted qnn03,

added qnn18, moved qnn33 to qnn03.  Reserved as101.  Deleted as105. Changed usage of

ptn11,acm04, and akn05.  Added hst table. Added ptn23.

06-22-01 Modified terminal identifier definitions in jag55, ins01, and unn02.  Added t_68a.

07-06-01 Added bd070.

07-10-01 Deleted unn18 & dfu18. Modified jag66 & jag67 & qnn16.  Modfied pt119, qnn31, dfq31.

Added hst13.   Deleted ptn23.

07-24-01 Updated bit definitions for aan03.

07-25-01 Rearranged qnn19 through qnn23 & dfq19 through dfq23. Expanded bit definitions for qnn31.

Updated commands in akn04.

07-31-01  Added jag02 and t_68b.

    - Added as101 definition.  Added some more status to arn05 and arn06.

    - Modified definition of  “AA” table.

    - Defined jag03. Redefined as102, as104, as105.  Added ro101-ro124. Added fsn14.

Added hst14-hst25 & changed length of hst records.

08-30-01 Corrected typo errors to jag02 and jag03.  Redefined usage of fwn16 and fwn19.

Assigned jag67 as terminal download status.

09-06-01 Changed “^” to “~” in akn10. Changed data type for acs06, acs07, akn06, arn06, arn07, & ptn12.

09-19-01 Add command value 9 to fsn04.

10-12-01 Changed ins01 character 3 usage.

10-15-01 Changed meaning of command code 11 for akn04 from “Display Message on Q.iMPACT” to

“Complete Feed in Manual Mode”.

11-14-01 jag02 is now a spare and is replaced by as108, jag03 is now a spare and is replaced by as109,

and jag59 is now a spare and is replaced by as110.  Swapped error codes 16 & 17 in arn06

11-30-01 jag02 has number of currently active PAC feeds. Added Power Failure Status to arn06 & arn07.

Added additional status bits to qnn31.  Changed APC to PAC.  Changed JAGXTREME to

Q.iMPACT.

01-07-02 Added last material transfer statuses 25 and 26 to arn06.

01-22-02 Made jag02, asn06 & asn07 spares.  Jag03 contains active PAC processes.  Added additional

status to arn05.

    - Added back in bd085 and bd086 which had been accidentally deleted.

    - Added cs157 as future run time counter and labelled cs257 – cs557 as spares.  Changed jag12 to

a spare.  Spares are no longer numbered.  Removed p_110 to p_127 (Physical Discrete Inputs 13 – 36),  p_510 to p_527 (Physical Discrete Outputs 13 – 36), p_6ec to p_6ef (Discrete Input rising edge 13 –16), and p_6fc to p_6ff (Discrete input falling edge 13 – 16).

    - Added jag59 for routing of data to Console Print Server (future addition to Q.iMPACT).

    - Added bd056 and bd057 and assigned csn60 (future addition to Q.iMPACT).

    - Corrected attribute on csn18 from 8 to 18.

    - Added bd024, bd025, and bd026 (future addition to Q.iMPACT).

Corrected format of spn01 from 8C to 9C.

Corrected format of jag55 from 3C to C and re-defined its use as flow meter ID only.

Re-defined use of jag02 as the selected flow meter location (local or remote) and is now used in conjunction with jag55.