# I.  Recording Changes

In order to comply with FDA CFR 21 Part 11 regulations and to meet security concerns expressed by P&G, the Q.iMPACT  has been enhanced to record changes to selected shared data variables.   The log file “sdwrites.log” (also referred to as the “Setup Changes” log) records each shared data write by storing the shared data variable name, date, time, author, and new value.  This file has a maximum size of 100,000 bytes.  An example printout of this log is shown later in this document.  Operator initiated changes are always recorded, however, the user has the option to record shared data writes initiated by the operating system itself as well as those initiated from FTP downloads, JagBASIC, or remote network shared data writes.  The following page, accessed from Maintenance – Logging Of Setup Changes, shows how to record these other shared data writes:

As more options are selected, the “sdwrites.log” file will fill more quickly.   This is especially true for  “System Initiated Changes”.  A warning message will be displayed once the log is at least 80% full and will be re-displayed every 10 minutes until the log file is saved and cleared.  Once the log file reaches 90% of capacity, a warning message will be displayed every 5 minutes.

The “sdwrites.log” file needs to be copied periodically to a  secure PC or a print-out made.  The log file must then be cleared.   The following web page, accessed from Maintenance – Setup Changes Log, shows how this is done:

The user can either dump the log file to the screen and then print it or  use FTP to transfer it to a PC.  Once this is done successfully, the log file must be cleared.

II.  Forcing an Operator Name / ID entry

All web page initiated shared data writes are recorded in the “sdwrites.log” file.  A shared data write record can include an author or can be anonymous.  In the case that the author must be recorded, a supervisor can force the entry of an Operator Name / ID prior to saving any changes or entering into critical areas.  The supervisor first needs to create a password using the MATROLLER Environment web page.  Once this is done, a key symbol will appear next to all buttons or links that require an Operator Name / ID entry.  Below is just one example of several web pages that require an Operator Name / ID entry to either save changes or access critical areas:

Notice the key symbol next to the Scale, Flow Meter, and PAC selections.  These groups require an Operator Name / ID entry  to save all changes made within the group or access critical areas, such as calibration.  Other groups such as MATROLLER or Maintenance only have one or two operations within the group that require an Operator Name / ID entry and are therefore not marked with the key symbol, however, the specific operation within the group is marked with a key symbol.

To disable forced Operator Name / ID entry, the supervisor must click on the Q.iMPACT terminal graphic (just below the MT Homepage button).   A separate window will appear that will permit the supervisor to disable forced Operator Name / ID entry.

III.  Recorded Shared Data Variables

ABC01 PLC address

ABC06 PLC data format

ABC85 ControlNet Media Layer

AS101 Purchased PAC channels

AS102 Purchased PAC material-paths

AS103 PAC database master

AS110 PAC weight entry method

CEn01 Address of First Load Cell

CEn02 Number of Load Cells

CEn03 Primary Units

CEn04 Primary Number of Ranges

CEn05 Primary Low Increment Size

CEn06 Primary Mid Increment Size

CEn07 Primary High Increment Size

CEn08 Primary LowMid Threshold

CEn09 Primary MidHigh Threshold

CEn10 Primary Scale Capacity

CEn11 Secondary Units

CEn12 Secondary Number of Ranges

CEn13 Secondary Low Increment Size

CEn14 Secondary Mid Increment Size

CEn15 Secondary High Increment Size

CEn16 Secondary LowMid Threshold

CEn17 Secondary MidHigh Threshold

CEn18 Secondary Scale Capacity

CEn19 Calibration Units

CEn25 Gravity Adjust

CEn26 Motion Sensitivity

CEn27 Motion Time Period

CEn85 Linearity Correction Enable

CEn88 Shift Adjust Mode

CS118 A scale ID

CS218 B scale ID

CS261 Operator Name / ID

CS318 C scale ID

CS418 D scale ID

CSn01 Auxiliary Units

CSn03 Custom Units

CSn04 Rate Time Units

CSn05 Rate Sample Time

CSn06 Rate Display Interval

CSn08 IDNET 10 Times Precision

CSn09 Power Up Timer

CSn10 Low Pass Filter

CSn11 Notch Filter

CSn12 Comb Filter

CSn17 Filter Poles

CSn19 Averaging Filter Order

CSn20 Comb Filter Order

CSn21 Scale Type

CSn22 Scale Location

CSn30 Calibration Date

CSn31 Next Calibration Date

CSn32 Calibration Days Interval

CSn33 Calibration Weightment Interval

CSn34 Calibration Expired Announcement Type

CSn36 Calibration Check Tolerance

CSn37 Calibration Check Points

CSn38 Calibration Check Failed Announcement Type

CSn40 Monitor Cell Overloads

CSn41 Cell Overload Threshold

CSn42 Monitor Platform Overload

CSn44 Monitor High Impact

CSn45 Impact Weight Threshold

CSn46 Impact Rate Threshold

CSn47 Monitor Weighments

CSn48 Weighment Trigger Type

CSn49 Weighment Threshold

CSn50 Monitor Zero Commands

CSn51 Monitor Zero Failures

CSn52 Monitor Scale IO Errors

CSn53 Weighment Reset Threshold

CSn54 Cell Overload Units

CSn55 Specified Run Flat Cell

CSn56 Percent Capacity Symmetry Start

CSn62 Cell Symmetry Check

CSn63 Cell Zero Drift Check

CSn64 Cell Zero Drift Threshold

CSn65 Cell Symmetry

CSn66 Cell Symmetry Threshold

CSn67 Predictive Fail. Announcement

CSn68 Run Flat Weight

CSn85 Fillnoise Filter Enable

CSn88 Rate Display Enable

CSn89 Auxiliary Weight Display Enable

CSn90 Units Switching Enable

CSn92 Do IDNET Tare In Q.i

CSn93 Process Application Type

DFQ01 Material-Path Assigned

DFQ02 Instrument Channel

DFQ03 Material-Path Descriptor

DFQ05 PAC Algorithm

DFQ06 Feeder Type

DFQ07 Slow Step Timer Factor

DFQ09 Minimum Open Time

DFQ10 Drain Time

DFQ11 Average Flow Rate Low Limit

DFQ12 Average Flow Rate High Limit

DFQ13 Average Spill Low Limit

DFQ14 Average Spill High Limit

DFQ15 Algorithm Update Parameter

DFQ16 Flow Rate Filter Sample Period

DFQ17 Destination Channel

DFQ18 Maximum Flow Rate Alarm Value

DFU01 Measuring Device Assigned

DFU02 Q.iMPACT Terminal

DFU03 Measuring Device

DFU04 Stable Scale Time

DFU06 Feed Override Time

DFU07 Minimum Slow Step Time

DFU09 Overlap Feed Alone Time

DFU10 Overlap Time Tolerance

DFU11 Zero Flow Threshold

DFU12 Unstable Device Flow Threshold

DFU13 Minimum Add

DFU14 Maximum Unit Size

DFU15 Dump Trip Point

DFU16 Minimum PAC Flow Rate Threshold

EMn01 Email Address

EMn02 Email Alert Messages

EMn03 Email Predictive Failures

FSn01 Flow Meter Description

FSn02 Flow Meter Increment Size

FSn03 Flow Meter Units

FSn07 Flow Meter Rate Units

FSn08 Flow Meter K Factor

FSn09 Flow Meter K Units

FSn11 Flow Meter M Factor

FWn03 Flow Meter Units

FWn07 Flow Meter Rate Units

INS01 Instrument Location (channels 1 – 20)

INS02 Instrument Location (channels 21 – 40)

INS03 Instrument Location (channels 41 – 60)

INS04 Instrument Location (channels 61 – 80)

INS05 Instrument Location (channels 81 – 100)

INS06 Instrument Location (channels 101 – 120)

INS07 Instrument Location (channels 121 – 140)

INS08 Instrument Location (channels 141 – 160)

INS09 Instrument Location (channels 161 – 180)

INS10 Instrument Location (channels 181 – 200)

JAG04 Market Region

JAG15 Number Of Internal Scales

JAG18 Consecutive Number Destination

JAG22 Consecutive Number Preset

JAG29 Monitor Log Records

JAG34 Q.iMPACT User ID

JAG35 Q.iMPACT Project,

JAG36 Q.iMPACT Description

JAG65 Battery Replacement Date

JAG88 Legal For Trade

JAG89 Consecutive Number Enable

JAG90 Consecutive Number Preset

LG801 ControlNet IP Address

LG802 ControlNet Subnet Mask

LG901 ControlNet Gateway Address

LG902 ControlNet DHCP Client Enable

LGn01 FTP User Name

LGn02 FTP User Password

NDE11 Cluster IP Address 1

NDE12 Cluster IP Address 2

NDE13 Cluster IP Address 3

NDE14 Cluster IP Address 4

NDE15 Cluster IP Address 5

NDE16 Cluster IP Address 6

NDE17 Cluster IP Address 7

NDE18 Cluster IP Address 8

NDE19 Cluster IP Address 9

NDE20 Cluster IP Address 10

NDE21 Cluster IP Address 11

NDE22 Cluster IP Address 12

NDE23 Cluster IP Address 13

NDE24 Cluster IP Address 14

NDE25 Cluster IP Address 15

NDE26 Cluster IP Address 16

NDE27 Cluster IP Address 17

NDE28 Cluster IP Address 18

NDE29 Cluster IP Address 19

NDE30 Cluster IP Address 20

NDE31 Host IP Address 1

NDE32 Host IP Address 2

NDE33 Host IP Address 3

NET02 Ethernet IP Address

NET03 Ethernet Subnet Mask

NET04 Ethernet Gateway IP Address

NET05 DNS Server 1 IP Address

NET06 DNS Server 2 IP Address

NET07 DNS Server 3 IP Address

NET10 Node Number Of Q.iMPACT

NET23 SMTP Server IP Address

NET24 Email Machine Name

NET25 Sender Email Address

NET26 Email Subject

NET28 SMTP Server TCP Port

NET30 Web Server Home Page

NET31 Web Server Document Page

NET32 Web Server Help Page

NET87 Enable Web Server

NET88 Enable FTP Server

NET89 Enable Email Alert

NET91 Enable Network Console

Qnn00 Material-Path Table Block

RM101 Node 1 Enabled

RM102 Node 2 Enabled

RM103 Node 3 Enabled

RM104 Node 4 Enabled

RM105 Node 5 Enabled

RM106 Node 6 Enabled

RM107 Node 7 Enabled

RM108 Node 8 Enabled

RM109 Node 9 Enabled

RM110 Node 10 Enabled

RM111 Node 11 Enabled

RM112 Node 12 Enabled

RM113 Node 13 Enabled

RM114 Node 14 Enabled

RM115 Node 15 Enabled

RM116 Node 16 Enabled

RM117 Node 17 Enabled

RM118 Node 18 Enabled

RM119 Node 19 Enabled

RM120 Node 20 Enabled

RM121 Host 1 Enabled

RM122 Host 2 Enabled

RM123 Host 3 Enabled

TRn01 Autotare Threshold

TRn02 Autotare Reset

TRn03 Autotare Clear

TRn85 Tare Enabled

TRn86 Pushbutton Tare Enabled

TRn87 Keyboard Tare Enabled

TRn88 Autotare Enabled

TRn89 Autotare Motion

TRn90 Autoclear Tare

TRn91 Autoclear After Print

TRn92 Autoclear Motion

TRn93 Tare Interlock Enable

TRn94 Display Tare

TRn96 Net Sign Correction

Unn00 Measuring Device Channel Table Block

ZRn01 Power Up Zero +

ZRn02 Power Up Zero -

ZRn03 Pushbutton Zero +

ZRn04 Pushbutton Zero -

ZRn05 Autozero Window

ZRn06 Under Zero Divisions

ZRn87 Autozero Gross-Net

ZRn88 Zero Gross

ZRn89 Zero Gross-Net

# IV.  Example Printout of Setup Changes Log – sdwrites.log

Below is a very short printout of the sdwrites.log.  It contains the five character shared data variable name, followed by the date, time, and author of the change.  The next line contains the new value.  An anonymous entry would have “- - - -“ (four dashes) for the author.

CS261 27 Sep 2002 09:49:29 William Smith

William Smith

FS102 27 Sep 2002 09:49:29 William Smith

0.01

FS103 27 Sep 2002 09:49:29 William Smith

2

FS107 27 Sep 2002 09:49:29 William Smith

S

FS108 27 Sep 2002 09:49:29 William Smith

2400

FS109 27 Sep 2002 09:49:29 William Smith

4

# V.  How it all works

The current shared data structure does not contain an element (or the room to add an element) that can be used as a flag to control logging of a write to a shared data variable.  Fortunately the root process used to create the shared data structure contains an unused security bit for each shared data variable.  By setting this bit, the shared data variable is marked for logging every time it is written.  The outputted structure does not have a security bit, so the upper bit of the length field is used to mark for logging.  This required modifying the operating system at several points to mask off this bit whenever the field length was used.  The Q.iMPACT has several entry points for saving shared data writes.  This presented a major problem that required modifying two lower level functions for recording a write operation.   Since there was no way to write to a RAM disk based file within a shared data write function, another solution needed to be found.  The solution involved creating two small ring buffers that hold the shared data variable index number or block number and the author type (‘O’ – operator entry, ‘S’ – system entry, ‘R’ – remote network entry , ‘F’ – FTP download, ‘J’ – JagBASIC entry).  Another function was created to read the ring buffers and convert the index or block number to a shared data name and store this along with the data, time, author, and new value in a RAM disk file.  This function is called periodically within the control panel thread.  In the case of a web page initiated write the control panel thread is awoken immediately after the write operation.  Shared data writes from the front panel are logged very quickly since a display update awakens the control panel thread, thus calling the logging function.  The operator name / ID (author) is always the first shared data variable (CS261) to be posted from the web page.  This guarantees that any following shared data writes posted contain the correct author.  Special comparison functions were also written to prevent logging shared data writes that did not change in value.  This is common to web page posting since a group of shared data variables is always posted.  Only the ones that change in value are logged.  Another change to the operating system required creating new shared data write functions that contained the author type so that changes made from the control panel or from other threads (such as JagBASIC) could be logged with the correct author.

# VI.  Additional Security

The MATROLLER Environment web page has been enhanced with a selection that allows the supervisor to disable the MATROLLER keypad.  This locks the keypad to prevent unauthorized tampering.  If a key is pressed on the keypad, the message “KEYPAD DISABLED” is displayed on the lower display for one second.  Note: Cycling power on the MATROLLER will re-enable the keypad.