Material Path Parameter defaults

For the actual description of the parameters see the Technical manual Chapter 4, subsection: assigning channels. Some parameters have recommended defaults (see below), although they will more than likely not be the most efficient settings this is a list of defaults that will get you feeding product successfully. These settings should be tuned by a specialist after initial feeds have been completed to help maximize overall system performance.

Material Transfer Parameters for Measuring Device: x

Max Flow Rate Alarm Threshold		0	(Disabled)

Feeder Type  -  GIW				You may select algorithm types  Spill Only or K1 to begin with

Feeder Type  -  LIW				You may select algorithm types  DTE, Spill Only or K1 to begin with

Process Times

Slow Step timer Factor			1.5

Minimum Open Time			2

Drain Time					6

Slow Feed Time (QiLite Only)		8

Alarm Only on Slow Step Timeout			this box must be checked

Average Flow Rate Limits

Low						0

High						2.5 times the flow rate you are expecting

Average Spill Limits

Low						halve the High Average Spill Limit below, and place a negative sign in front

High						2 to 3 times the spill  you are expecting

Other Parameters

Algorithm Update				Spill Only use 0.5,    K1 or K2 use 0.75

Flow Rate sample period			2

Destination Channel		255 if outside the Cluster else the channel that this Material is feeding into