CODES D'ERREUR

| 1.0 | Transient error(hardware error, probably static discharge).scale needs service |
| --- | --- |
| 1.1 | Balance temperature transducer hardware error |
| 2.0 | Balance is unable to stabilize within time limit after tarring. Environment is too hostile or balance needs recalibration. |
| 3.0 | Incorrect or no calibration mass used for calibration. Recalibrate with correct masses. |
| 3.1 | Balance was unable to acquire stable data during internal calibration. The balance will try again. |
| 4.4 | RS232 buffer is full. |
| 7.0 | User entry out of bounds |
| 7.2 | Number outside of display capacity. |
| 8.0 | Hardware error causing an internal weight signal, which is too low. Check if pan is off. If not, balance needs service. |
| 8.1 | Hardware error caused by an internal weight signal, which is too high. Check load on the pan, which may be excessive. Persisting, balance needs service. |
| 8.2 | Power-on load out of specification. |
| 8.3 | Rated capacity exceeded. Removed excessive weight from pan. |
| 8.4 | Underload condition on balance. Check that the proper pan is installed. |
| 8.5 | Internal calibration weight internal sensor indicated its weight on the pan. |
| 9.1 | Bad factory checksum. If error persists, have the balance serviced. |
| 9.2 | Bad factory checksum. If error persists, have the balance serviced. |
| 9.3 | Bad factory checksum. If error persists, have the balance serviced. |
| 9.4 | Factory internal calibration data failed checksum. This failure will disable access to the InCal feature. |
| 9.5 | Factory calibration data failed checksum. |
| 9.8 | User calibration data failed checksum. |
| 9.9 | Factory temperature compensation data failed checksum. |