CODES D'ERREUR

| ERREUR | CONDITION | DESCRIPTION |
| --- | --- | --- |
| n n n n n n | OVERLOAD | Weigh on scale is greater than calibrated capacity range. |
| u u u u u u | UNDERLOAD | Displayed weight is less than calibrated capacity range. |
| id iLL | ID ILLEGAL | Id number 0 cannot be stored. Id number must be a positive integer. |
| NO CAL | NO CALIBRATION | Configurations settings relating to scale capacity, divisions, grads, or decimal point have been changed. Recalibrated scale using current settings |
| id inu | NO TOLERANCE | No tolerances have been entered for the displayed id register |
| buffer | BUFFER FILLING | Buffer 90% full |
| full | BUFFER FULL | Buffer 100% full |
| fail s | FAIL SIGNAL | No load cell signal present. |
| do err | DATA OUTPUT ERROR | Cannot be lb/oz data in 6 display digits. |
| Err cl | CALIBRATION ERROR | CAL SP value must be greater than 20% of full capacity. |
| Err ls | LOW SIGNAL ERROR | Less than 0.3uv/grad signal is present. |
| btn X | KEYPAD BUTTON FAILURE | Keypad button x is not producing a signal |
| NO UNI | NO UNITS | EPROM checksum failure. CW-80 has default to originals units. |
| id Err | ID ERROR | EPROM write failure. Unable to write to id selected. |