( Rice Lake )    IQ 310 CALIBRATION PROCEDURES

TO CALIBRATE VIA THE FRONT PANEL KEYBOARD, PROCEED AS FOLLOWS:

1. SET THE OPERATE/SETUP SWITCH ON THE REAR PANEL TO THE SETUP POSITION.

2. SCROLL RIGHT WITH THE PRINT KEY UNTIL CALIBRATE IS DISPLAYED.

3. SCROLL DOWN ONCE WITH THE DOWN ARROW (GROSS/NET) KEY TO DISPLAY ZERO.  IF ZERO DOES NOT APPEAR ON THE DISPLAY, SCROLL RIGHT UNTIL ZERO IS DISPLAYED.
4. CLEAR THE PLATFORM OF ALL WEIGHT.

5. PRESS TARE (ENTER).  THE MESSAGE *CALIB* APPEARS ON THE DISPLAY, INDICATING THAT IT IS CALIBRATING THE DEADLOAD.

6. WHEN THE DEADLOAD CALIBRATION IS FINISHED, SPAN APPEARS ON THE DISPLAY INDICATING THAT THE INDICATOR IS READY TO CALIBRATE THE SPAN.

7. PRESS THE DOWN ARROW KEY ONCE.  THE NUMBER DISPLAYED WILL BE THE FULL SCALE.

8. IF THE CERTIFIED TESTING WEIGHT EQUALS FULL SCALE CAPACITY, SKIP TO STEP 9.  IF THE TEST WEIGHT IS LESS THAN FULL SCALE CAPACITY, THE NUMBER ON THE DISPLAY MUST BE ADJUSTED TO MATCH THE TEST WEIGHT VALUE.  THIS CAN BE DONE BY USING THE PANEL KEYS.  NOTE THAT IF THE DECIMAL POINT LOCATION HAS BEEN PROGRAMMED IN THE DEC PNT PARAMETER, THE TEST WEIGHT VALUE MUST BE ENTERED EXACTLY AS IT WILL APPEAR WITH THE DECIMAL POINT WHEN WEIGHING.

9. PRESS TARE .  THE MESSAGE *CALIB* APPEARS ON THE DISPLAY INDICATING THAT THE SPAN IS BEING CALIBRATED.  (IF THE DISPLAYED NUMBER REMAINS AND *CALIB* DOES NOT APPEAR, IT INDICATES YOU HAVE TRIED TO ENTER A SPAN VALUE THAT IS OUTSIDE THE SCALE AS DEFINED BY THE FIRST FOUR PARAMETERS OF THE SETUP MENU).

10. AFTER THE SPAN CALIBRATION IS FINISHED, TW ZERO APPEARS ON THE DISPLAY.

11. PLACE THE OPERATE/SETUP SWITCH BACK IN THE OPERATE POSITION AND INSTALL THE SWITCH COVER AND LEGAL SEAL.

