# ( Rice Lake )    IQ 5000 Calibration Procedures

- IF YOUR SCALE HAS A REMOTE SCALE OPTION, CHECK AGAIN TO ENSURE THAT IT IS IN THE “LOCAL” MODE.

- CHECK TO BE SURE THAT THE SCALE BASE IS “EMPTY” (NO LOAD), AND THAT THE PLATFORM IS ON AND SEATED CORRECTLY.

- CLOSE (UPWARD) SWITCH #3 (“DEAD LOAD”).  THE MSD (LEFTMOST) DISPLAY DIGIT SHOULD BE FLASHING “C”.

- ADJUST THE POTENTIOMETER MARKED “DEAD LOAD” TO OBTAIN A READING AT OR NEAR “ZERO” (0).  NOTE THAT A READING OF EXACTLY “0” IS NOT NECESSARY.  ALSO NOTE THAT A CLOCKWISE ADJUSTMENT INCREASES (MAKES MORE POSITIVE) THE WEIGHT READING.

- OPEN (DOWNWARD) SWITCH #3.

- CLOSE SWITCH #1.  THE WEIGHT MAY CHANGE FROM A “0” READING, WHICH IS ACCEPTABLE.

- PUSH THE CLEAR KEY.  THE DISPLAY SHOULD READ “CAL” AND THEN RETURN TO A “0” READING.

- NOW PLACE WEIGHT ON THE SCALE OR LOAD SCALE CORRESPONDING TO AT LEAST 10% OF FULL SCALE CAPACITY.

- THE SPAN MAY BE CORRECTED IN EITHER OF TWO WAYS:

- PUSH THE “ZERO” KEY TO “ROLL UP” THE SPAN OR PUSH THE “DISPLAY MODE” KEY TO “ROLL DOWN” THE SPAN.  NOTE THAT AFTER A FEW SECONDS THE SPAN STARTS TO “ROLL” FASTER.

- OR -

- ENTER THE SPAN VALUE SEQUENTIALLY VIA THE KEYBOARD.

- AFTER THE DISPLAY READING HAS REACHED THE CORRECT VALUE, PUSH THE “SAMPLE” KEY.  THE DISPLAY WILL READ “CAL” FOR A SECOND, AND THEN RETURN TO THE CORRECT WEIGHT READING.

- RECHECK THE “ZERO” WEIGHT BY EMPTYING THE SCALE.

- OPEN SWITCH #1.

- REPLACE THE CENTER BAR AND TIGHTEN THE SCREWS.