# ( Rice Lake )    IQ 600 & 700 CALIBRATION PROCEDURES

WITH NO LOAD ON SCALE:

- SET DEAD LOAD SWITCH ON AND ADJUST DEAD LOAD POTENTIOMETER TO NEAR ZERO.

- APPLY LOAD GREATER THAN 50%.

- USE KEYPAD TO ENTER DESIRED READING AND PUSH “ENTER”.

- RETURN SWITCH TO OPERATE MODE.

NOTE:

- AZM OFF DURING CALIBRATION

- DEAD LOAD DISPLAY PROVIDES SIGNAL REF.

- STRAIGHT AVERAGING APPLIED DURING CAL.