( SETRA )                LINEARITY CALIBRATION

Linearity Calibration Procedure for Quick Count

Locate the CAL/PROTECT switch

To find the switch without removing the scale cover, look inside the hole, underneath the front

panel. Using a screwdriver, or allen wrench tool provided by Setra, push the switch to the left

(Cal position). If you are unable to locate the switch, power down the scale and remove the scale

cover. The CAL/PROTECT switch is on the bottom edge of the main board, near the middle of

board and go to step 2.

To make sure the switch is in the CAL position with the power on, press the ZERO key. The display

should read PROTECT.

- Power down the scale.

- Depress and hold the SAMPLE SIZE key and PRINT key simultaneously and power the scale up. Release both keys.

- The display should read D-NOS.

- Press the CLEAR key. The display will read BUSY, then PASS.

- Press the CLEAR key again. The display will read CAL FS.

- Press the SAMPLE SIZE key until the scales full scale capacity appears on the display. ( the readouts is in grams).

- When the scales full capacity is displayed, press the COUNT key. The display will read BUSY, then CAL A.

- Press the COUNT key. The display read LOAD 5.

- Load the full scale weight onto the pan.

- Press the COUNT key. The display will read BUSY, then LOAD 2. Remove 60% of the weight from the pan.

- Press the COUNT key. The display will read BUSY, then LOAD 0. Remove all the weight from the pan.

- Press the COUNT key. The display will read BUSY, then LOAD 1. Load 20% of full scale onto the pan.

- Press the COUNT key. The display will read BUSY, then LOAD 2. Load 40% of full scale.

- Press the COUNT key. The display will read BUSY, then LOAD 3. Load 60% of full scale.

- Press the COUNT key. The display will read BUSY, then LOAD 4. Load 80% of full scale.

- Press the COUNT key. The display will read BUSY, then LOAD 5. Load full scale on the pan.

- Press the COUNT key. The display will read BUSY, then PROTECT.

- Push the CAL/PROTECT switch to the right (PROTECT position).

- Press the clear key. The total weight reading on the pan is displayed in grams. The QUICK COUNT scale is now calibrated.

NOTE: After linearity calibration, the scale will display in extended resolution

(One order Greater of magnitude). To return to normal operation, simply

power down the scale.