# ( SETRA)               QUICK COUNT CALIBRATION PROCEDURES

- REMOVE ALL OBJECTS FROM THE PAN.

- LEVEL THE SCALE.

- PRESS THE UNITS KEY TO PLACE THE SCALE IN EITHER THE GRAMS OR POUNDS WEIGHING MODE.

- PRESS THE ZERO KEY AND WHILE THE SCALE IS DISPLAYING “busy”, PRESS THE UNITS KEY.  THE SCALE WILL FLASH “CAL” AND RETURN TO DISPLAYING “busy”.

- THE SCALE WILL THEN DISPLAY “L xxxxx”, WHERE xxxxx IS THE SCALE’S FIRST CALIBRATION POINT.  PRESS THE SAMPLE SIZE KEY REPEATEDLY, THE SCALE WILL DISPLAY EACH CALIBRATION POINT AVAILABLE.

- WHEN THE DESIRED CALIBRATION POINT IS DISPLAYED, LOAD THE DISPLAYED WEIGHT ON THE PAN, AND PRESS THE UNITS KEY.

- THE SCALE WILL DISPLAY “busy”, AND THEN RETURN TO THE NORMAL WEIGHING MODE.

- REMOVE THE WEIGHTS FROM THE PAN, PRESS ZERO TO ZERO THE DISPLAY.