Trooble Shooting

DF-1500

| TROOBLE SHOOTING GUIDE |  |
| --- | --- |
| SYMPTOM | CAUSE/SOLUTION |
| the indicator will not turn on no display | 1. no power- check AC power source. 2. check power connection. 3. defective power supply. replace. |
| prog error on power-up. | 1. bad indicator board- requires service. |
| setup error on power-up. | 1. setup parameters not saved. re-calibrate. 2. parameters corrupted - requires services |
| under on display | 1. load cells not wired correctly. 2. defective load cells. replace |
| over on display | 1. scale is overloaded. 2. dead load or span not set properly- re-calibrate 3. load cell not wired correctly. 4. defective load cell. replace. |
| lovolt on display | 1. low voltage power input- check AC power source. 2. defective power supply. replace. |
| Zero on display and cannot zero scale | 1. weight is outside zero windows- empty or clean scale. 2. incorrect deadload setting- reset deadload. |
| display unstable | 1. loadcell cabling wet or damaged- dry or re-wire. 2. live load not set correctly. 3. poor grounding- check ground wiring. |
| no output to printer, host computer or scoreboard. | 1. interface cable disconnected. reconnect. 2. serial interface not wired correctly. 3. output port not set up correctly. set parameters. |