ERROR MESSAGES

10	Cannot soan a negative value.

11	Span table full.

20	Parameter memory write error. indicator needs services.

21	Parameter checksum error. parameters have been lost.

22	Program check fault. indicator needs services.

34	Load cell input value out of range.(8-35mv).

41 	Pushbutton tare is invalid(over/under/motion).

44	Cannot do dead load calculation, input signal is greater than parameter 11.

45 	Cannot do dead load calculation, input signal is too far negative.

uuuuuu	loadcell not wired correctly. defective loadcell.deadload has not been set.

eeeeee		scale is overloaded. overweight setting is not correct. deadload or span not set properly. loadcells not wired correctly. defective loadcell. unused scale channel disabled.

rrrrrr	deadload has not been set. weight on scale is beyong zero window. scale has not been calibrated.