ALLEGANY

CRANE SCALE

Sur la façade enlever la petite plaque sous le bouton TARE

Potentiomètre du haut pour le Span

Potentiomètre du bas pour le Zero

| # Message | # Signification |
| --- | --- |
| Error | The display did not return to zero (no load) after a totalization operation. Remove bad from scale and allow display to return to zero before attempting the next totalization operation. Error will also appear if an incorrect entry is attempted. i.e. Calibration attempted while scale is in operation mode. An attempt to calibrate over capacity of scale. An attempt to zero over 100% of Scale Capacity. |
| bAnd | Appears during motion tolerance and zero track procedures as a prompt for the operator to enter the band width. |
| SPAn | Indicates that span factor(s) have been initialized |
| SP-Err | Proton Weight Processor is spanned to a count greater than 50. Linearity and breakpoints are reset. Span factor Is reset to one. Reset linearity, if necessary. Enter a span factor of 50 or less. |