CODE D'ERREUR

transcell ti-1500

| CODE | MODE | MEANING/POSSIBLE SOLUTION |
| --- | --- | --- |
| ooooooo | normal operating mode | gross overload. a weight greater than the rated capacity has been applied to the scale. remove the weight from tha platter or try re-calibrating the scale. otherwise, check for a bad loadcell connection or possible loadcell damage due to overloading. |
| Err 0 | span calibration mode(F17) | keyed-in value is less than 1% of full scale capacity. use a larger test weight or check keyed-in value. |
| Err 1 | span calibration mode(F17) | keyed-in value is larger than full scale capacity. use a smaller test weight or check keyed-in value. |
| Err 2 | span calibration mode(F17) | there is not enough load cell signal to produce the internal counts necessary to properly calibrate the scale. first check all load connections. use F16 to view internal counts. |
| Err 3 | all modes | non-volatile memory read error. one or more parameters have been lost. |
| Err 4 | all modes | non-volatile memory write error. indicator needs service. |
| Err 5 | key-in span calibration mode(F20) | you have attempted to enter a zero value for C1. enter a known calibration value greater than zero. |
| Err 7 | initialization | no reading from the ADC. make sure there is a load cell connected to the indicator at start-up. |
| Err 9 | normal operating mode | span calibration value has been lost. re-calibrate the scale. |
| Err 10 | initialization or setting code entry | indicator cannot read serial number. indicator needs service. |
| Err 11 | initialization | serial number mismatch. re-enter setting code from product ID tag |
| Err 12 | initialization or setting code entry | invalid setting code. re-enter proper setting code from ID tag |
| Err 13 | setting code entry | non- volatile memory is in wrtie-protect mode. toggle position of setup/calibration switch or JP2 jumper. |