| # Interfacing a Mettler Toledo Balance Home Newsletters Catalogue Search Top of Form Bottom of Form | Windmill software lets you interface Mettler Toledo balances, and other instruments with an RS232 serial port, to a PC running Windows. Our new Windmill COMIML software is very easy to use and you can try before you buy. Alternatively, subscribers to the Monitor newsletter (ISSN 1472-0221) can download an earlier version of the driver (LabIML) for free. This page helps you get started with the free Windmill 4.3 and LabIML software. (Configuration is similar with COMIML, but you use the comDebug program instead of the ConfIML program). There are three basic steps. - Run the Windmill ConfIML program and enter your balance's communication settings. - Run the Windmill SetupIML program and choose names and units for your balance's data. - Run the Windmill DDE Panel, Logger or Chart program to save or display data from your balance. After you have entered your configuration settings with ConfIML and SetupIML, you don't need to use these again and can go straight to the logging and display programs. More details on entering the configuration settings are below. ## Run the Windmill ConfIML program This detects and saves a record of the hardware you want to use. You need to add the LabIML ASCII instrument handler and enter your instrument's communication settings. To do this - Press the Add button. - Select LabIML RS232 ASCII Instrument Handler. - Type a name for the balance and a description. The name and description can be anything you like. You will generally enter 1 as the number of channels. (Each channel holds one type of data item: in this case weight is the only value being measured.) Press OK. - Enter your instrument's settings. These should work for Mettler Toledo balances: - Reading Protocol: Request/Response On Demand - Timeout: 5000 - Instrument Idle Time: 0 - Count how many characters are in the balance's message - you should find this information in the balance's Data Interface Operating Instructions. If not set to 20. - Instrument Initialisation String: Leave blank. - You now need to define how to extract data. To do this press the Channels button. Your settings here depend upon how you have configured your balance. Enter a command or prompt string to send to the balance to ask for data, then enter parsing instructions to extract the data value from the balance's reply. Check your balance's Interface Operating Instructions for details of the commands to send, and the format of the balance's reply. For example, to request a single stable value you would set... - Prompt String: S followed by a Carriage Return and a Line Feed. This looks like: S\C013\C010 - Parse String: Ignore 3 characters, Extract until space.. This looks like: \I" "\E"\ " - The next step is to enter your communications settings. These should be the same as those used by the balance: check the Mettler "Data Interface Operating Instructions" for your model. If in doubt start with 9600 baud, 8 data bits, no parity, 1 stop bit and no flow control. - Save your settings and close ConfIML. ## Now Run the Windmill SetupIML program Here you can name your channel, set its units, any alarms and so on. - From the Device menu select LabIML. - Your data channel will be shown as a number like 10000. Double click this channel. - Type name for your channel and make sure Enable for Input is checked. If you want to change the units enter an appropriate scale and offset. For example, if you wanted to convert milligrams to grams you'd enter 1000 as the scale and 0 as the offset. - Save your settings in a *.ims file and close SetupIML. ## Now open Windmill DDE Panel - From the File menu select Load Hardware Setup and choose the *.ims file you just saved. - Connect your weight channel. You should see the correct weight in DDE Panel. - Proceed similarly for the Logger and Chart programs. ## Getting the data into Excel You can use the Windmill Logger program to collect data, and after collection has finished import it into Excel. Alternatively, you can collect data with Excel in real-time by writing an Excel macro to read data from the Windmill DDE Panel. For more details see our Excel page. ## Success? Have you successfully used Windmill with your balance? Let us know about your project. Top of Form Details of how you are using Windmill: Name: Organisation: E-mail: Bottom of Form |
| --- | --- |