Flash the Software

The first step is to extract the new Jaguar files from the floppy disk onto the personal computer. To extract the files:

Do not perform the file download from within the Windows environment. Exit Windows and perform the following steps from the DOS prompt.

Create a directory and go to the directory on the PC where the new files will be stored. Use the DOS MD command if you need to create a new directory to hold the Jaguar files, and the CD command to change directories.

- Insert the diskette with the new software files into the floppy disk drive A or B. The files on this diskette are compressed into a single file named JAGUAR.EXE.

- At the DOS prompt, type A:JAGUAR or B:JAGUAR depending on where you inserted the diskette, then press ENTER.

The compressed files will automatically extract themselves from JAGUAR.EXE into the designated directory on the computer. The extracted files require approximately 2.6 MEG of storage space on the PC’s hard drive.

- Edit the FP.BAT file using the DOS text editor or another and verify the new software name and serial port are correct. The FP.BAT file resembles the example below. Edit only the -t and -com elements.

Example:

flashpro -tSYS9_16.HEX -b115.2 -com1 -pe -d7 -s1

where:
-t is followed by the file name to be saved such as SYS9_16.HEX.

-b is the baud rate (in kbaud). The software is transferred at 115.2 kbaud. DO NOT CHANGE THIS PARAMETER.

-com is the serial com port on the personal computer that will be used to transfer the new software to the Jaguar terminal.

-p is the parity. The parity is set to even. DO NOT CHANGE THIS PARAMETER.

-d is the number of data bits used. Data bits are set to seven. DO NOT CHANGE THIS PARAMETER.

-s is the number of stop bits used. Stop bits are set to one. DO NOT CHANGE THIS PARAMETER.

You are now ready to use the flashpro batch file (FP.BAT) to load the software into the Jaguar’s flash memory.

To load Jaguar software:

If a DigiTOL base is installed, be sure to disconnect all wires prior to connecting the RS-232 cable.

Disconnect power to the Jaguar.

- Remove the Controller board and remove the W2 (WDI) jumper. Reinstall the Controller board without the jumper.

- Connect a bi-directional RS-232 cable from a personal computer to Jaguar’s COM2 serial port. Wire the cable as follows:

| Jaguar COM2 | 9-pin COM | 25-pin COM |
| --- | --- | --- |
| TxD B RxD B Gnd Jumpers | 2 3 5 4 to 6, 7 to 8 | 3 2 7 4 to 5, 6 to 20 |

- With the cursor in the directory containing the new Jaguar files, type FP at the DOS prompt, but DO NOT press ENTER to execute the command.

- Using a small, non-conducting object such as a bare Q-Tip, press and hold the pushbutton switch behind the CAL opening on the back of the Controller PCB.

- With the CAL switch still depressed, apply power to the Jaguar then press ENTER or RETURN on your computer immediately to execute the flashpro installation program.

- Release the CAL switch when a framed box scrolling a string of A’s appears on the computer monitor indicating the installation has begun.

The box will continue to scroll A’s until the installation procedure is complete. This process will continue with brief pauses occasionally for about 10 to 15 minutes. If the box with A’s does not appear, press ESCAPE on the computer and remove power from the Jaguar. Repeat steps 4 through 7.

If the new software was downloaded successfully, DOS prompt will be restored on the computer and the Jaguar will perform a power-up sequence.

During the power-up sequence, the storage locations for setup parameters and memory fields are checked. If any have been moved or added in the new version of software, you must respond Yes to the following prompts:

BRAM Bad - Rst? Y or BRAM Err - Rst? Y

EError - Reset? Y or EE A Err - Rst? Y

- Press ENTER to reset the Jaguar’s program block parameters to factory defaults. If you select N(o), you will encounter scale errors.

If BRAM Err - Rst? Y is displayed, you must reconfigure the Jaguar.

If EError - Reset is displayed, you must recalibrate the scale.

If EE A Err - Rst? Y is displayed, you must recalibrate the scale.

- Disconnect power from the Jaguar then remove the Controller board and reinstall the W2 (WDI) jumper. Finally, reinstall the Controller board with the jumper in place.