ADDENDUM TO LYNX® INDUSTRIAL TERMINAL SERVICE MANUAL

The following features and enhancements have been added to the Lynx terminal containing the newly released revision “C” software.  This software should be present in any Lynx terminals with a July 1997 (-4UY) or later date code.  The software version can be confirmed by viewing it during the powerup sequence.  Please add this information to your Lynx service manual (part number 14467800A) at the locations described below.

The following information is an update to page 3-1 (Operation Mode column).  The Password sub-block was not shown.

- The following information is an addition to pages 3-48 and 3-49.

Password Configuration

Operator access to setpoint values can now be password protected.  There is only one password in the system for all activities.

- The following information is an update to page 3-11 (step 10).

Gravity Adjust

The values for gravity adjustment have been changed to a new series of numbers in the range 0 to 31.  The operator is now prompted:  [ Geo Val?XX ].  Refer to the following chart for the Geo Value if used.

| Geo Value Table | Height above sea level in meters |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 0 325 | 325 650 | 650 975 | 975 1300 | 1300 1625 | 1625 1950 | 1950 2275 | 2275 2600 | 2600 2925 | 2925 3250 | 3250 3575 |
| North or south latitude | Height above sea level in feet |  |  |  |  |  |  |  |  |  |  |
| in degrees and minutes | 0 1060 | 1060 2130 | 2130 3200 | 3200 4260 | 4260 5330 | 5330 6400 | 6400 7460 | 7460 8530 | 8530 9600 | 9600 10660 | 10660 11730 |
| 00 0’ - 50 46’ | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 1 | 1 | 0 | 0 |
| 50 46’ - 90 52’ | 5 | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 1 | 1 | 0 |
| 90 52’ - 120 44’ | 6 | 5 | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 1 | 1 |
| 120 44’ - 150 6’ | 6 | 6 | 5 | 5 | 4 | 4 | 3 | 3 | 2 | 2 | 1 |
| 150 6’ - 170 10’ | 7 | 6 | 6 | 5 | 5 | 4 | 4 | 3 | 3 | 2 | 2 |
| 170 10’ - 190 2’ | 7 | 7 | 6 | 6 | 5 | 5 | 4 | 4 | 3 | 3 | 2 |
| 190 2’ - 200 45’ | 8 | 7 | 7 | 6 | 6 | 5 | 5 | 4 | 4 | 3 | 3 |
| 200 45’ - 220 22’ | 8 | 8 | 7 | 7 | 6 | 6 | 5 | 5 | 4 | 4 | 3 |
| 220 22’ - 230 54’ | 9 | 8 | 8 | 7 | 7 | 6 | 6 | 5 | 5 | 4 | 4 |
| 230 54’ - 250 21’ | 9 | 9 | 8 | 8 | 7 | 7 | 6 | 6 | 5 | 5 | 4 |
| 250 21’ - 260 45’ | 10 | 9 | 9 | 8 | 8 | 7 | 7 | 6 | 6 | 5 | 5 |
| 260 45’ - 280 6’ | 10 | 10 | 9 | 9 | 8 | 8 | 7 | 7 | 6 | 6 | 5 |
| 280 6’ - 290 25’ | 11 | 10 | 10 | 9 | 9 | 8 | 8 | 7 | 7 | 6 | 6 |
| 290 25’ - 300 41’ | 11 | 11 | 10 | 10 | 9 | 9 | 8 | 8 | 7 | 7 | 6 |
| 300 41’ - 310 56’ | 12 | 11 | 11 | 10 | 10 | 9 | 9 | 8 | 8 | 7 | 7 |
| 310 56’ - 330 9’ | 12 | 12 | 11 | 11 | 10 | 10 | 9 | 9 | 8 | 8 | 7 |
| 330 9’ - 340 21’ | 13 | 12 | 12 | 11 | 11 | 10 | 10 | 9 | 9 | 8 | 8 |
| 340 21’ - 350 31’ | 13 | 13 | 12 | 12 | 11 | 11 | 10 | 10 | 9 | 9 | 8 |
| 350 31’ - 360 41’ | 14 | 13 | 13 | 12 | 12 | 11 | 11 | 10 | 10 | 9 | 9 |
| 360 41’ - 370 50’ | 14 | 14 | 13 | 13 | 12 | 12 | 11 | 11 | 10 | 10 | 9 |
| 370 50’ - 380 58’ | 15 | 14 | 14 | 13 | 13 | 12 | 12 | 11 | 11 | 10 | 10 |
| 380 58’ - 400 5’ | 15 | 15 | 14 | 14 | 13 | 13 | 12 | 12 | 11 | 11 | 10 |
| 400 5’ - 410 12’ | 16 | 15 | 15 | 14 | 14 | 13 | 13 | 12 | 12 | 11 | 11 |
| 410 12’ - 420 19’ | 16 | 16 | 15 | 15 | 14 | 14 | 13 | 13 | 12 | 12 | 11 |
| 420 19’ - 430 26’ | 17 | 16 | 16 | 15 | 15 | 14 | 14 | 13 | 13 | 12 | 12 |
| 430 26’ - 440 32’ | 17 | 17 | 16 | 16 | 15 | 15 | 14 | 14 | 13 | 13 | 12 |
| 440 32’ - 450 38’ | 18 | 17 | 17 | 16 | 16 | 15 | 15 | 14 | 14 | 13 | 13 |
| 450 38’ - 460 45’ | 18 | 18 | 17 | 17 | 16 | 16 | 15 | 15 | 14 | 14 | 13 |
| 460 45’ - 470 51’ | 19 | 18 | 18 | 17 | 17 | 16 | 16 | 15 | 15 | 14 | 14 |
| 470 51’ - 480 58’ | 19 | 19 | 18 | 18 | 17 | 17 | 16 | 16 | 15 | 15 | 14 |
| 480 58’ - 500 6’ | 20 | 19 | 19 | 18 | 18 | 17 | 17 | 16 | 16 | 15 | 15 |
| 500 6’ - 510 13’ | 20 | 20 | 19 | 19 | 18 | 18 | 17 | 17 | 16 | 16 | 15 |
| 510 13’ - 520 22’ | 21 | 20 | 20 | 19 | 19 | 18 | 18 | 17 | 17 | 16 | 16 |
| 520 22’ - 530 31’ | 21 | 21 | 20 | 20 | 19 | 19 | 18 | 18 | 17 | 17 | 16 |
| 530 31’ - 540 41’ | 22 | 21 | 21 | 20 | 20 | 19 | 19 | 18 | 18 | 17 | 17 |
| 540 41’ - 550 52’ | 22 | 22 | 21 | 21 | 20 | 20 | 19 | 19 | 18 | 18 | 17 |
| 550 52’ - 570 4’ | 23 | 22 | 22 | 21 | 21 | 20 | 20 | 19 | 19 | 18 | 18 |
| 570 4’ - 580 17’ | 23 | 23 | 22 | 22 | 21 | 21 | 20 | 20 | 19 | 19 | 18 |
| 580 17’ - 590 32’ | 24 | 23 | 23 | 22 | 22 | 21 | 21 | 20 | 20 | 19 | 19 |
| 590 32’ - 600 49’ | 24 | 24 | 23 | 23 | 22 | 22 | 21 | 21 | 20 | 20 | 19 |
| 600 49’ - 620 9’ | 25 | 24 | 24 | 23 | 23 | 22 | 22 | 21 | 21 | 20 | 20 |
| 620 9’ - 630 30’ | 25 | 25 | 24 | 24 | 23 | 23 | 22 | 22 | 21 | 21 | 20 |
| 630 30’ - 640 55’ | 26 | 25 | 25 | 24 | 24 | 23 | 23 | 22 | 22 | 21 | 21 |
| 640 55’ - 660 24’ | 26 | 26 | 25 | 25 | 24 | 24 | 23 | 23 | 22 | 22 | 21 |
| 660 24’ - 670 57’ | 27 | 26 | 26 | 25 | 25 | 24 | 24 | 23 | 23 | 22 | 22 |
| 670 57’ - 690 35’ | 27 | 27 | 26 | 26 | 25 | 25 | 24 | 24 | 23 | 23 | 22 |
| 690 35’ - 710 21’ | 28 | 27 | 27 | 26 | 26 | 25 | 25 | 24 | 24 | 23 | 23 |
| 710 21’ - 730 16’ | 28 | 28 | 27 | 27 | 26 | 26 | 25 | 25 | 24 | 24 | 23 |
| 730 16’ - 750 24’ | 29 | 28 | 28 | 27 | 27 | 26 | 26 | 25 | 25 | 24 | 24 |
| 750 24’ - 770 52’ | 29 | 29 | 28 | 28 | 27 | 27 | 26 | 26 | 25 | 25 | 24 |
| 770 52’ - 800 56’ | 30 | 29 | 29 | 28 | 28 | 27 | 27 | 26 | 26 | 25 | 25 |
| 800 56’ - 850 45’ | 30 | 30 | 29 | 29 | 28 | 28 | 27 | 27 | 26 | 26 | 25 |
| 850 45’ - 900 0’ | 31 | 30 | 30 | 29 | 29 | 28 | 28 | 27 | 27 | 26 | 26 |

- The following information is an update to pages 3-16 (figure) and 3-17 (step 10).  Net sign correction has been moved from the demand connection section (page 3-26) to the Tare block.

3333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333T333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333cleared of all data fields and left blank33333333333333333333333333333333333333333333333333333333in response to the 8142 333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333

3333

3333

3333

| 3333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333 |  |  |  |  |
| --- | --- | --- | --- | --- |

333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333333344444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444444nction from a host.  The “Save” function baud rate is limited by that of the other device connected to COM1.  On conclusion of the save or load operation, the display shows a completion message and waits for the operator to acknowledge by pressing any key.   The number shown on the right is the number of setup fields successfully transferred.  The Save/Load data format is ASCII text which can be transferred by the DOS *.Bat files shown below which uses the file name LYNX.TXT for the setup file.  Any commercially available data communication program running under DOS or Windows that supports a “text file” protocol may also be used.

3333

3333

3333333

3333

µ §

Save.Bat

MODE COM1: BAUD=2400 PARITY=E DATA=7 STOP=1 RETRY=P

COPY COM1: LYNX.TXT

Load.Bat

MODE COM1: BAUD=2400 PARITY=E DATA=7 STOP=1 RETRY=P

COPY LYNX.TXT COM1:

- The following information is an update to page 4-15.

PRIVATE  §Storing a Temporary ID/Tare Recordtc "5.6.3.6.1   Process ID/Tare"§

During the Temporary ID/Tare store process an Auto Assigned ID will now briefly display the assigned ID so that the operator may record the record number for future reference.

- The following information is an update to page 4-24.

Function key access to ID/Taretc "5.6.3.7.4   Access ID/Tare Table"§ data

While accessing a Single ID/Tare you will now be prompted “Another?Y” at the end of the access to expedite multiple ID/Tare entries.

µ §

- The following information is an update to page A-17.  This code was not used previously.

|  | Function | Code |  | Data Field |
| --- | --- | --- | --- | --- |
| Function Code Description | Hex | ASCII | Direction | Length |
| Last Printed Data from Printer Port | 51 | Q | Upload | 45 |

- The following is new information to add to page A-29.

Last Printed Data (Upload)

This function code will return the data that is currently in print template ptp03.  If ptp03 is cleared to a blank condition, the response from the Lynx duplicates the format of the 8142 instrument.  The 8142 format is shown below.

| Host Transmission |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| ASCII | STX | 2 | U | Q | CR |
| Hex | 02 | 32 | 55 | 51 | 0D |

| Lynx Response |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| STX | 2 | U | Q | Data from ptp03 | CR |
| 02 | 32 | 55 | 51 | (default 45 bytes) | 0D |

Format for ptp03 when it is cleared.

| Data | Length |
| --- | --- |
| Gross Weight Tare Weight Net Weight Time Date CN (ID)Var01 | 7 7 7 6 6 6 6 |

If ptp03 is not cleared out, the Lynx operates in a slightly different mode.  It uses approximatley 4K of memory as a transaction buffer to store all fields programmed in ptp03 for each transaction.  To retrieve the transaction information, the host sends the “Q” command to receive the oldest record then an “H” command to clear this first record.  The host then sends another “Q” command for the next transaction then an “H” to clear that record.  Using this sequence, all transactions can be retrieved from the Lynx.  The following is an example of this mode.

ptp03

- The following information is an update to page A-40 (SICS version number).

The Lynx  SICS command set has been updated from version 2.0x to version 2.10.

- The following information is an update to page A-46 (item 4.).

4.  I4 - INQUIRY OF SERIAL NUMBER

The response to this command will include any data programmed into Literal 8 (Lit08).  If you wish the Lynx to return the actual serial number of the unit, the serial number must be manually  programmed into literal 8.

- The following information is an update to page A-47 (item 6.).

There was an error in the response to the SI command in earlier versions of software.  The incorrect response included an extra letter “S”.  This command now works as described in the manual

- The following information is an update to page A-49 (item 10.).

There was an error in the operation of the @ command in earlier versions of software.  This command now works as described in the manual.  Note that  any tare value will be cleared and the Lynx will respond with the I4 response (I4_A_"4380720-4TY" ).

- The following information is an update to page A-52 (item 3.).

There was an error in the operation of the TAC command in earlier versions of software.  The command required to clear tare was actually “TA_C”.  The command now works as described in the manual.

Windows ‘95 Hyperterminal Instructions

Setup

- Access the “HyperTerminal” program group (“Start”, “Programs”, “Accessories”, “HyperTerminal”)

- Run the HyperTerminal program by clicking on the “Hypertrm” icon.

- Enter a New Connection name (“SaveLoad” for example) and choose an icon then click “OK” to proceed to the “Phone Number” window.

- Change the “Connect using” to “Direct to Com X” where X is the port you are using then click on “OK” to proceed to the “COMX Properties” window.

- Select the baud rate (see the opening paragraph above), data bits, and parity desired.  Set “Stop bits” to “1” and “Flow control” to “None”.

- Click on “Advanced” then click on the 666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666666 close the “COMX Properties” window.

- Click on “File”, “Properties” and “Settings”.

- Set “Emulation” to “TTY” then click on “ASCII Setup”.

- Click on the “Send line ends with line feeds”  box to add a check mark.

- Click on the “Wrap lines that exceed terminal width” box to remove the check mark.

- Click on “OK” to return to the main properties box.

- Click on “OK” to exit the main properties box.

Saving setup data from the Lynx

- Click on “Transfer”, then “Capture Text”.

- Change the “File” setting to the desired path and filename and then click on “Start”.

- Enter Lynx setup confirm that the COM1 baud rate, data bits, and parity setting match those selected on the PC.

- Proceed to the “Save Setup” prompt and press Enter on the Lynx.

- The “Saving...” prompt  on the Lynx should be displayed for approximately 30 seconds before a “Saved 2xx”. display is shown.

- Click on “Transfer”, “Capture Text”, “Stop”.

- Note that you must delete the existing file if you wish to save new data in a file that already exists.  (Hyperterminal concatenates data onto an existing file).

Loading setup data to the Lynx

- Click “Transfer”, “Send Text File” and enter the path and filename in the “Filename” box (or use the icons to navigate to the desired file location).  Do not click “Open” at this time!

- Enter Lynx setup (confirm COM1 baud rate, data bits and parity setting if not done previously) then proceed to the “Load Setup” prompt and press Enter on the Lynx.

- Click on “Open” on the PC screen within 10 seconds.

- The “Loading...” prompt on the Lynx should be displayed for approximately 30 seconds before a “Loaded 2xx” display is shown.